Manual

RK2675 series(upgraded version) Leakage current tester

### The highest output voltage of this series tester can run up to 250V

###### * Any incorrect operation may cause death.

* Please read the second chapter"the use of precautions"before using the tester.

* The operating instruction should be placed next to the operator so that you can read it when you need it.

Instructions for use

###### * Please read and understand the content described in the instructions

###### carefully before operating tester. After reading,please put the manual

###### near to operator for reading when needed. When the tester was carried

###### from a workplace to another place.Please carry the instructions with

###### the instrument to prevent loss.

* If it is found that the page is missing or the Instructions is contaminated.Please contact the dealer of the Meiruike  for purchase immediately.

* With the improvement of the function of the instrument, upgrade of the software.The instructions will also continue to be improved,upgraded.

Please note that the software of test equipment and manual version.

RK2675 series(upgraded version) Leakage current tester manual Ver:1.0	2016.08

In order to ensure safety, please note.

###### * In the testing process,the operator should not touch the location or area of the following;otherwise it will cause the occurrence of electric shock accident.

（1）High voltage output port of test instrument;

（2）The test line connecting with the tester;

（3）Tested product;

（4）Any object that is connected to the output port of the tester.

* In order to prevent the accident of electric shock,please follow the steps below：

（1）Grounding connect safely and reliablely:There is a grounding terminal on the rear faceplate of this series of tester,please connect this terminal to the ground.If there is no reliable grounding connection,it is likely to cause the electric shock as long as any human contact the shell.So it must ensure the terminal connectted with earth reliablely.

（2）After turn on the power switch,please don’t touch any objects connected with the high voltage output port.

* The following situation is very dangerous:

（1）After pressing the "STOP" key,the test indicator light is still on;

（2）The voltage value of the display don’t changed any more and the test indicator is still on;

Meet the above situation,turn off the power switch and unplug the power plug immediately,don’t use it any more;please contact with the Meiruike instrument or designated dealer immediately.

CONTENTS

Unpacking and Installation of Tester	5

1.1  Receiving instrument examination	6

1.2  Unpacking inspection of instrument	6

1.3  Packing box and packing material	6

1.4  Safety rules for the use of instruments	6

1.5  Operator requirements	7

1.6  Check voltage of power supply	7

1.7  Check and replace the fuse	8

1.8  Grounding connection	8

Operating considerations	9

2.1  Forbidden operation	10

2.2  Emergency handling	10

2.3  Preventive measures in the testing process	10

Description of faceplate	11

3.1  Description of front faceplate	12

3.2  Description of rear faceplate	13

Summary of instrument function	14

4.1  Summary	15

4.2 Introduction of instrument function	15

Technical Specifications	16

5.1 Comparison table of type function	17

5.2 Technical parameter	17

Interface of PLC	18

6.1  Input,output signal of PLC interface	19

6.2  Wiring	19

6.3  Connection instructions of remote control input signal and output signal	19

6.4  Electrical characteristics of PLC interface	19

Setting of instrument parameter	20

7.1  Setting of RK2675AM parameter	21

Instrument test function	22

8.1  Test specification	23

8.2  Test step	23

8.3 Test chart which GB9706.1 prescribed	24

Accessories and warranty	27

9.1 Accessories	27

9.2 Warranty	27

# 1

## Unpacking and Installation of Tester

This chapter describes the basic condition of the test and installation of the tester after user received it.

##### 1.1 Receiving instrument examination

After you receive the test instrument of the Meiruike instrument, according to the following steps to check:

1.1.1 Whether the packing box of instrument is intact;if damaged,we recommend that you do not carry out of the box,and connect with the Meiruike company or the distributor of Meiruike.

1.1.2  If the packing box of instrument is well, please check whether the type of the instrument you ordered and the type on the packing box is consistent,if no,please connect with Meiruike company or the distributor of Meiruike.

If after 1.1.1 and 1.1.2 examination,there are no problem,you can check out of the box of the instrument.

##### 1.2 Unpacking inspection of instrument

Please check whether the attachment and the 9.1 attachment list in the package you ordered is consistent.

##### 1.3 Packing box and packing material

Please keep the original packing material in order to use it in the transportation.

##### 1.4 Safety rules for the use of instruments

When using the instrument,please follow the safety rules:

###### 1.4.1 Don’t use test equipment in flammable air.

In order to prevent theexplosion or burning accident occurr,do not use the tester in the presence of alcohol,thinner or other combustible materials,and do not use the instrument in the air containing flammable gases.

###### 1.4.2 Don’t use the test instrument in the area of high temperature or direct sunlight.

The components used in the instrument are precision parts,which should be avoided used in the high temperature or direct sunlight place,it will accelerate the aging of the instrument,shorten the life of the tester,it is also possible to damage the tester.

Using temperature range of instrument:0ºC to 40ºC;

Storage temperature range of instrument:-20ºC to 70ºC;

1.4.3 Don't use tester in a high humidity environment.

Don't put the instrument on the boiler,kettle,humidifier or water in high humidity environment.Congealed droplet may cause internal short circuit and damage the tester, seriously may cause fire.If the environment humidity of storage instrument exceeds specified below,the test must be dry completely before use.

Using humidity range:20%～80%RH

Storage humidity range:less than 90%RH

###### 1.4.4 Don’t use the tester in a dusty environment

A dusty environment may cause a short circuit in the instrument to cause a fire.

###### 1.4.5 Don’t use the tester in poor ventilated environment

Forced air cooling system in internal instrument;it will overheat and damage the instrument if the heat inside the instrument doesn’t go out.Don’t stack other objects in the air inlet and outlet of the instrument to block the air duct.

###### 1.4.6 Don’t put the tester on the inclined surface or use the tester in the shaking of the place.

Instruments placed on a sloping or shaking surface will make the tester fall off easily.

###### 1.4.7 Don’t use the tester in the sensitive test equipment or beside the receiving equipment.

If used the tester beside these devices,these devices may be interfered by the tester;in order to reduce the interference of these devices,so that these devices should be away from the tester.

###### 1.4.8 The input power of the tester must have a separate switch control

The input power of the tester must have a separate switch control,once the emergency situation occured should be cut off the power switch to deal with the accident immediately.

##### 1.5 Operator requirements

The output voltage of the tester is sufficient to cause death,so it is necessary to be qualified personnel to operate the test instrument.

###### 1.5.1 Personnel qualification

It must be operated by a skilled person to understand the basic concepts of voltage, current and resistance;The operator must know that high pressure is flowing from the high voltage output port of tester when in high pressure tester,pass the tested body,it flows into the tester through the connection line of the current return port.

###### 1.5.2 Safety rules

Operators must undergo special training,and understand all kinds of safety rules and safety regulations,read the instructions carefully before operating the tester.

###### 1.5.3 Regulations of dress

Operatiors should not wear clothes with metal decorative or wear metal accessories, such as watches and so on.

##### 1.6 Check voltage of power supply

###### 1.6.1 Switching input supply voltage of instrument

This instrument use the 220V AC±10% 50Hz±2Hz single-phase power supply,before opening the power switch on front faceplate of instrument,please ensure that the power supply voltage and fuse is consistent with the voltage of the instrument faceplate which selected by switch.

###### Warning:in order to prevent failure or damage to the tester, please use the test instrument in the specified

###### voltage range.

##### 1.7 Check and replace the fuse

Warning:

* In order to avoid the accident of electric shock, before the replacement of the fuse, please turn off the power switch and unplug the power plug.

* Make sure that the fuse used is in accordance with the shape, size and characteristics specified in the specification of the tester.Otherwise it may damage the tester.

1.7.1 Fuse specifications

| Type | Fuse specifications |
| --- | --- |
| 2675AM | 5A |

1.7.2 Fuse replacement

When replacing the fuse,please follow these steps:

（1）Turn off the power switch on the front faceplate and pull out the power cord;

（2）Use a screwdriver to fix the fuse holder as shown below:

（3）Check fuse specifications and replace the fuse to the same specifications as the fuse listed in 1.7.1;

（4）Put the supporting frame back to the original position.

##### 1.8 Grounding connection

###### Warning:* There may be an accident of electrical breakdown grounding connection in the wrong place or not connecting.

In order to ensure the safety,we must ensure that the instrument is grounding connection reliablely;There are two ways to ensure reliable grounding connection of the instrument,please select one to connect the instruments with ground the reliablely.

（1）Connect the power cord to a power socket of three-phase grounding;

（2）If the three-phase power supply socket is not grounded,there is a protective grounding terminal on the rear faceplate of the instrument,and connected it to the safety ground.

# 2

## Operating considerations

This chapter describes the norms, measures and precautions must be followed by the operating instruments;please read the contents of this chapter in detail before using the instruments;

Warning:It is necessary to be very careful and follow the warnings given in this chapter,notes and other instructions when operate the instrument.

##### 2.1 Forbidden operation

###### 2.1.1 It is forbidden to continuous and rapid switching the power supply switch

Turn off the power switch on the front panel,if you want to re open the power switch, please ensure that turn off the power switch in time a few seconds or longer.Don’t switch the power switch repeat and frequently,if so,protection apparatus may not be appropriate to implement the protection function;Please do not turn off the power switch when the tester is testing the output voltage,unless it can be executed in an emergency.

###### 2.1.2 It is forbidden to touch the measured object during testing

It is charged for leakage current measurement,the measured shell is charged.So operators must pay attention to safety,constitute the relative safety operating procedures.Don’t touch the device to be tested before cut off the power supply,in order to prevent electric shock,the specific grounding methods please refer to 1.8.

##### 2.2 Emergency handling

In case of emergency(the occurrence of an electric shock or the test object combustion),you must take the following actions;you can complete (1) or (2) first,but both must be completed.

（1）Closing the power switch of instrument;

（2）Unplug the power cable from the power socket.

##### 2.3 Preventive measures in the testing process

###### 2.3.1 Prevent electric shock proof with insulated gloves

In order to prevent the electric shock accident,please bring the rubber gloves with insulation and then engage in electrical related work before using this tester.

###### 2.3.2 Check whether the instrument can work normally

After the instrument is connected to the power supply,press the open button of the power supply to observe whether there is abnormal situation in the instrument faceplate,once appear,turn off the power immediately.If display normal,without access to the tested product.Check the start,reset,N/L switch,automatic/manual,time setting button,whether it works properly.Once you find an unusual situation,that stop testing, and please contact the Meiruike or distributors.

###### 2.3.3 The tested object are connected to the tester  correctly

When connecting with the measured object,the output voltage must be "0" and "reset"

（1）When testing the device I,insert the plug of the measured object into the tester's three socket,press the test button at the top left.

（2）When testing the device Ⅱ,insert the plug of the measured object into the tester's binocular socket,press down the Ⅱ switch on the lower left.

###### 2.3.4  Replace the measured object

When a test object has been tested and replace another object,please confirm:

（1）The tester is in a "reset" state.

（2）The test indicator doesn’t twinkle.

###### 2.3.5 Tester is in test condition

When the tester is in the test condition,the test line,the test object,the output port are all equipped with high pressure,please do not touch.

###### 2.3.6 Stop testing

Please turn the power switch placed in the location of OFF when don’t use the tester any more,or you need to leave.

# 3

## Description of faceplate

### This chapter describes components of the front faceplate and rear faceplate of the tester,please know all the functions on the instrument faceplate before operating the instrument.

##### 3.1 Description of front faceplate

##### 3.2 Description of rear faceplate

3.1 Description of front faceplate

###### 3.1.1	Description of RK2675AM front faceplate

The front faceplate of RK2675AM

1、Input power switch

Press for open (ON),pop-up for(OFF).

###### 2、START  test button

Press this button to start the test in the boot condition.

###### 3、STOP stop button

Press for the test indicator lamp off without test voltage output.

4、Leakage test voltage adjusting knob

Clockwise for large,counterclockwise for small.

###### 5、Output power socket

The output of the power outlet is divided into I type tester output and II type tester output.When using the II class tester output,the 6 button must be in the "II" state;

###### 6、I class and II class switch button

When the tested product belongs to I device,the button is in the off position.

7、II class measuring input port

When the tested product belongs to II device,this port is connected with the internal MD for the current measurement terminal.

8、Super leakage indicator lamp for leakage current

This lamp light indicates that the leakage current is out of range.

###### 9、Test status indicator lamp

This lamp light indicates that the instrument is in the test state.

###### 10、Voltage display window

11、Current display window

12、Time display window

The range of time is 0.0s~999s.When the time is less than 100s,the time resolution is 0.1s;when the time is greater than or equal to 100s,the time resolution is 1s.If the time is set to 0.0s,then the test time to add count;when the time is not set to 0,the time to reduce count.

13、UP button

When setting the time, press this key, the time setting value will increases;

14、DOWN button

When setting the time, press this key, the time setting value will decreases;

15、N/L  phase switch button

16、Auto/manual phase switch button

###### 17、Current preset adjusting potentiometer

When the "test/preset"(18)button is pressed for the current preset state,the current display window shows the preset current value,adjust clockwise the preset current will increases;adjust anticlockwise the preset current will decreases.

18、Preset、test button

This button press for preset,pop-up for test;

19、2mA/20mA  switch  button

This button is pressed for the 20mA file,pop-up for the  2mA file.

##### 3.2 Description of rear faceplate

###### 3.2.1 Description of RK2675AM  rear faceplate

1、Input power outlet

Three-core two-phase power supply socket.The power socket is equipped with a fuse;please refer to 1.7.about the selection of the input voltage and the replacement of the fuse.

2、Protecting ground port

This protecting grounding port must be connected to the protective ground reliably.Otherwise,the shell of the tester may be filled with high pressure,causing the occurrence of electric shock.

3、Interface of PLC

About the detailed description of the PLC interface,please refer to the sixth chapter.

# 4

## Summary of instrument function

### In this chapter, the function of the tester is described in detail. Please read the contents of this chapter before operating the tester.

##### 4.1 Summary

RK2675 series leakage tester is designed in accordance with IEC, ISO, BS, UL, JIS and other international and domestic safety standards.Leakage output voltage is 0~250V adjustable continuously.The maximum output power can reach 3000VA;It suitable for leakage current test of high electric system of all kinds of household appliances,power supply,all kinds of lamps,electrical,medical,chemical,electronics,instrumentation,machine.It is also a indispensable leakage current test equipment of scientific research laboratory and technical supervision departments.

RK2675AM satisfy the standard of GB4706.1-2005,GB/T12113-1996.

##### 4.2 Introduction of instrument function

###### 4.2.1	Automatic phase switching function of RK2675AM

N/L phase switching function can be selected automatically,when the instrument works in the automatic switching function,the instrument change phase every 2 seconds automatically.

###### 4.2.2	Timing test function

When setting the timing value(greater than 0)and testing,start timing then reaches the setted value,stop timing and cut-off voltage automatically.When the test time is set to "00.0s",the tester will continue to test until the user presses the "STOP" button to stop the test after starting test.

###### 4.2.3	RMS  demodulation

The voltage and current adopt the RMS demodulation,it can adapt to all kinds of different load test.

5

## Technical parameter

The technical parameters of various tester are introduced in detail in this chapter.

##### 5.1 Comparison table of type function

| Type | Function description |
| --- | --- |
| RK2675AM | Leakage current tester with maximum output voltage of 250 and maximum output power of 500W |

##### 5.2 Technical specifications

5.2.1 R K 2675AM

| Type | RK2675AM |
| --- | --- |
| Output voltage range | （0-250）V |
| Maximum(power) output | 500VA |
| Upper limit of current setting | 10uA~20.00mA |
| Test time setting | (0.0~999)s |
| Switch phase method | Automatic, manual |
| Voltmeter | Range | （0-250）V |
|  | Accuracy | ±3% |
|  | Resolution | 1V |
| Ammeter | Range | 10uA~20.00mA |
|  | Accuracy | ±5% |
|  | Resolution | 2mA files：1uA；20mA files：10uA |
| Interface of PLC | Optional |
| Standard | GB4706.1-2007 |

6

## Interface of PLC

The usage of PLC interface is described in this chapter.

The remote control terminal is attached on the back board of the tester,it can connect to remote operation.The terminal is standard 9PIN D type terminal block,it divide into input signal terminal and output connection terminal.

##### 6.1 Input,output signal of PLC interface

Resetting

##### 6.2 Wiring

TEST control:The control switch is connected between the PIN1 and the PIN3.

RESET  control:The control switch is connected between the PIN1 and the PIN4.

Testing signal output:Between PIN2 and PIN5.

Test failure signal:Between PIN6 and PIN7.

Test qualified signal: Between PIN8 and PIN9.

##### 6.3 Connection instructions of remote control input signal and output signal

The tester is equipped with a remote control point,which can be operated by the external remote control device of TEST and RESET functions.These contacts provide a power supply with a control function,the "momentary contact" switch must be used as a controller.

Special note:It can not be connected to any other power supply absolutely,if you input the other power will cause damage to the internal circuit of the instrument.

The output signal provides relay contact output.

##### 6.4	Electrical characteristics of PLC interface

Output contact withstand the maximum voltage:24V AC/DC;the maximum current 100mA

The input terminal is connected to the non-voltage control contact,null terminal voltage:＜10VDC

# 7

## Setting of instrument parameter

### The parameter setting of the tester are introduced in detail in this chapter.

##### 7.1  The parameter setting of RK2675AM

7.1 The parameter setting of RK2675AM

###### 7.1.1 Presetting of test time

7.1.1.1 The preset conditions of test time

The tester must be in a reset state,that is the tester can’t be in a state of test and alarm.

7.1.1.2 Preset method

（1） Time increases

In the front faceplate has a UP button,click on this button,the time preset value plus 1;if you hold this key,preset time continuous add 1,when added to a certain value,preset time continuous add 10;until 999s.

（2） Time decreases

In the front faceplate has a DOWN button,click on this button,the time preset value minus 1;if you hold this key,preset time continuous minus 1,when reduced to a certain value,preset time continuous minus 10;until 0.0s.

###### 7.1.2 Presetting of AC current alarm value

7.1.2.1 The preset conditions of AC current alarm value

Select the corresponding files,this series tester with 2mA, 20mA two files.

7.1.2.2 Preset method

Press the "test/preset" button,the current display window shows the current preset value;use the debug driver clockwise to adjust preset adjust potentiometer of current,preset current increased;adjust counterclockwise,preset current is reduced;set the alarm value of current to the required value of the test,pop up the "test / preset" button.

###### 7.1.3 Adjustment of output voltage

When the voltage regulator knob is in the 0 place,even if start the tester,it has no high voltage output;after starting the tester,adjust the output voltage adjusting knob clockwise,the output voltage increased,and reach the voltage required for testing.

# 8

## Instrument test function

### In this chapter, we introduce the testing function of all kinds of testing instruments.

##### 8.1 Test specification

###### 8.1.1 Human network

Human network of  RK2675AM：

###### 8.1.2 Equipment classification

（1）Class I equipment-electric-shock safeguard of equipment not only depends on the basic insulation,but also includes the additional safety measures,which connecting the accessible conductive parts to the protective conductor of the facilities fixed line. To make accessible conductive parts not charged when in the basic insulation failure.

（2）Class II equipment-electric-shock safeguard of equipment not only depends on the basic insulation,but also includes the additional safety measures.Such as double insulation or reinforced insulation,but it is no grounding connecting or dependent on the protective facilities of installation conditions.

###### 8.1.3 Equipment can be tested of RK2675A series

RK2675 series tester can test I class device or II class device.Through the button of the front faceplate to select the current measurement port.

##### 8.2 Test procedure

###### 8.2.1 Connection graph of RK2675AM

Connecting of RK2675AM and I class device

Connecting of RK2675AM and Ⅱ class device

8.2.2 Setting parameters

Please refer to the Chapter 7.

###### 8.2.3 Start testing

Press the start button to start testing,during the test,if the test current is less than the upper limit of current,then test end and the test lamp light off;if the test current is greater than the upper limit current,the super leakage indicator lights and buzzer.

##### 8.3 GB9706.1 test pattern

###### Figure 16 With or without the measurement circuit of application part I class device to ground leakage current

Figure 17 Use the specified class I single-phase power supply,with or without the measurement circuit of application part I class device to ground leakage current

Figure 18 Measuring device for the leakage current of shell

9

## Accessories and warranty

## 9.1 Accessories

1、Power cord	1 root

2、Instruction manual	1 copy

3、Quality assurance	1 copy

4、Product qualification certificate	1 copy

After the user receives the instrument,should check the contents of the box,if there is a shortage,please contact with our company or distribution unit.

##### 9.2 Warranty

9.2.1 Warranty Period (1)、The use units purchase instrument from the company,calculate since the date of shipment of the company,purchase from

the dstribution department,calculate since the date of shipment of the distribution unit,the whole machine warranty period of 12 months.

(2)、The warranty period of perishable products such as accessories is 6 months.

9.2.2 Warranty

It should issue the instrument warranty card when repair.Our company provide the lifelong maintenance services for all external equipment.In the warranty period,due to the improper operation of the user and damage to the instrument,the maintenance costs will assumed by theser.

The products of Meiruike company have been approved and are being examined and approved by the Chinese patent protection.

The information provided in this instructions replaces all the information that is published in the past.Our company reserves the right of change the specifications and prices.And won’t give any notice further.

Meiruike company reserved all rights

# Text boxes

Failure

Qualified

Testing

Test

5

4

3

2

1

9	8

7

6

Quaed

失败