SHENZHENG E MEIRUIKE C ELECTRONICY TECHNOLOGY CO.,LTD

RK9910/RK9920 Series

AC/DC Withstanding Voltage & Resistance Tester

Manual V1.0

CONTENTS

Chapter 1: Installation and use …………………………………………………………………………………1

1.1 Precautions for use …………………………………………………………………………………1

1.2 Points to note when moving………………………………………………………………………1

1.3 Connect AC Power Cord………………………………………………………………..2

1.4 Ground……………………………………………………………………………………2

1.5 Operation check………………………………………………………………………….3

1.6 Other characteristics of the instrument………………………………………………..3

Chapter 2: Codes of Practice and Measures…………………………………………………4

2.1 Prohibited Operation…………………………………………………………………….4

2.2 Handling of emergencies……………………………………………………………….4

2.3 Preventive measures during testing…………………………………………………..4

2.4 High Voltage Test Warning…………………………………………………………......5

2.5 Dangerous state treatment of faulty instruments…………………………………….6

2.6 Conditions to guarantee long-term trouble-free use…………………………………6

2.7 Daily inspection………………………………………………………………………..6-7

Chapter 3 Instrument Panel Overview…………………………………………………………7

3.1 Front panel description………………………………………………………………….7

3.1.1 Power Switch (POWER)…………………………………………………………..8

3.1.2 START key, STOP key……………………………………………………………. 8

3.1.3 Trademarks and models………………………………………………………….. 8

3.1.4 USB interface………………………………………………………………………8

3.1.5 Indicator area………………………………………………………………………8

3.1.6 FUNCTION…………………………………………………………………………8

3.1.7 HV……………………………………………………………………………………9

3.1.8 High-end output voltage (HV)……………………………………………………..9

3.1.9 Test low side, test current return side (LOW / RET)…………………………. 9

3.1.10 Move key………………………………………………………………………..…9

3.1.11 Shortcut Function Keys………………………………………….………………9

3.1.12 LOCK key………………………………………………………………………9

3.1.13 LCD………………………………………………………......………………9

3.1.14 USB Interface……………………………………………......……………………9

3.1.15 LOCK ……………………………………………......…………………………….9

3.2 Rear panel description………………………………………………………………….. 9

3.2.1 Amplifier Fan Cooling Port……………………………………………………….10

3.2.2 Test low end, test current return end (option) ………………………………..10

3.2.3 High-voltage output (option) …………………………………………………..10

3.2.4 Power socket and power selection switch…………………………………….10

3.2.5 Protective ground terminal………………………………………………………10

3.2.6 Nameplate……………………………………………......………………………10

3.2.7 RS232C……………………………………………......…………………………10

3.2.8 USB Device………………………………………......…………………………10

3.2.9 HANDLER interface……………………………......………………………….10

3.2.10 SINGLE interface……………………………......……………………………10

3.2.11 Interface………………………………………......……………………………11

3.2.12 Power Transfer Switch……………………......………………………………11

3.3 Instrument Performance Overview………………………………………………11-14

Chapter 4: Basic Operations…………………………………………………………………15

4.1 Overview of Instrument Interface Structure………………………………………..15

4.2 Panel function interface and parameter description……………………………..16

4.2.1 SETUP measurement settings. The interface is as follows……………..17

4.2.2 TEST test interface. The interface is shown as follows: (take AC as an example) ……………………………………......………………………………….…18

4.2.3 SYSTEM system interface. The interface diagram is as follows: …….19-20

4.2.4 FILE file storage interface………………......…………………………..….…21

4.3 Test project interface and parameter description………………………………….22

4.3.1 AC AC withstand voltage test parameter setting. The setting interface is as follows: …………………………………………………………………………………22

4.3.2 DC DC withstand voltage test parameter setting. The setting interface is as follows: …………………………………………………………………………………23

4.3.3 IR insulation resistance test parameter setting. The setting interface is as follows: (Schematic 4.3.3) ……………………………………………………………24

4.4 Test function principle and instructions………………………………………………25

4.4.1 Start test…………………………………………………………………………..25

4.4.2 Test delay………………………………………………………………………..25

4.4.3 Voltage rise……………………………………………………………………..26

4.4.4 DC boost determination………………………………………………………..26

4.4.5 High voltage test……………………………………………………………..26

4.4.6 Test voltage drop……………………………………………………………..26

4.4.7 Ground Current Detection Function………………………………………….26

4.4.8 Current overrun and arc detection (ARC) function……………………..….27

4.4.9 Unqualified Judgment………………………………………………………....28

4.4.10 Test result processing……………………………………………………....28

4.4.11 STOP (Stop Measurement) ……………………………………………....29

4.4.12 OFFSET…………………………………………………………………….29

4.5 HANDLER interface and SINGAL interface circuit structure and use………..30

4.5.1 Principles of Control Interface………………………………………………..30

4.5.2 Control Interface Use………………………………………………………..31

4.6 Other interfaces and functions of the instrument……………………………..31

Chapter 5 Explanation of Serial Instruction Set ……………………………………….31

Brief description of the instruction format: ……………………………………………..31

5.1 SCPI instruction set………………………………………………………………..31

5.2 DISP LAY subsystem command set…………………………………………..31

5.3 FUNC TION subsystem command set…………………………………………..32

5.3.1 FUNCtion subsystem command set is mainly used to set the test parameters of the instrument test function………………………………………………………………..32

5.3.2 PROG Function Command Set……………………………………………….33

5.3.3 AC Setup Function Command Set…………………………………………. 34-36

5.3.4 DC Setup Function Command Set …………………………………………36-38

5.3.5 IR Setup Function Command Set …………………………………………..38-39

5.4 SYSTEM Subsystem Command Set …………………………………………40-42

5.5 MMEM subsystem command set…………………………………………………43

5.6 FETCH Subsystem Command Set………………………………………………43

5.7 Other control command set………………………………………………….……44

6 Technical parameters ………………………………………………….…………….44

6.1 Specific parameters ………………………………………………….…………44-47

6.2 General technical indicators………………………………………….…………..47

Chapter 7: Explanation of Instrument System Upgrade Procedures ……….…………..48

Chapter 8: Appendix Parameter Table ……………………………………….…………….48

Chapter 9: Program-controlled medical pressure tester parameter table, random accessories, manual revision history…………………………………….………………...49

Chapter 10: Program-controlled medical pressure tester parameter table, random accessories, manual revision history………………………………….…………………....50

Chapter 1 Installation and Use

This chapter describes some of the checks that must be performed when you receive the instrument, and the conditions you must know and have before installing and using the instrument.

1.1 Precautions for use

The following rules must be observed when using the instrument:

■ Do not use the instrument in combustible air

To prevent burning or explosion, do not near alcohol, thinner and other flammable materials.

Use the instrument in air.

■ Avoid exposing the instrument to high temperatures and direct sunlight

Do not place the instrument in a place subject to heat or extreme changes in temperature.

Instrument operating temperature range: 5 ° C to + 35 ° C

Instrument storage temperature range; -20 ° C to + 60 ° C

■ Avoid humid environment

Do not place the instrument in a boiler, humidifier or high humidity environment with water.

Instrument humidity range: 20% to 80% RH (no dew condensation allowed)

Instrument storage humidity range; less than 90% RH (no dew condensation allowed)

Condensation may cause the circuit to malfunction. You must wait until the environment is completely dry before using the instrument.

■ Do not place the instrument in an environment with corrosive gases

Do not use the instrument in an environment with corrosive gases such as sulfuric acid, mist or the like. This may corrode the wires and connectors, forming

Hidden dangers or connection defects can cause malfunctions, failures or even fires.

■ Do not use the instrument in a dusty environment

Dirt and dust can cause electrical shorts or fires.

■ Do not use the instrument in a place with poor ventilation

The instrument has a forced air cooling system. Provide sufficient space for side and rear air vents to ensure air circulation.

■ Do not use the instrument on inclined surfaces or in places subject to shaking

If the instrument is placed on an uneven surface or a place where it is shaken, the instrument may slip and damage the instrument.

■ Do not use the instrument in places with strong magnetic or electric field effects

If the instrument is used in a strong magnetic or electric field, electromagnetic pulses may cause the instrument to malfunction and cause a fire.

■ Do not use the instrument near sensitive test equipment and receiving equipment

If these devices are used near this instrument, noise from the breakdown of the device under test may affect these devices. Measurements over 3kV

Test voltage. The electric field between the test leads will ionize the air and cause corona, which will cause a lot of RF (radio frequency) bandwidth interference between the test leads. for

In order to reduce this effect, ensure that the distance between the test leads is sufficiently long.

Also, keep test leads away from conductive surfaces (especially sharp metal ends).

1.2 Points to note when moving

When moving the instrument or transporting it, pay attention to the following precautions:

■ Before moving, turn off the power switch

Moving with the power switch on can cause electric shock and damage.

■ Disconnect all cables before moving

Moving the instrument without disconnecting the cables may result in damage to the cables or the instrument may fall during handling.

1

1.3 Connect AC Power Cord

The power cord is provided by the company with the instrument. Do not use an AC power cord that is not supplied with the instrument.

Connection order

1． Make sure the power supply is within the line power range of the instrument.

2． Determine the nominal value of the instrument fuse, and the fuse box is installed in the correct position (power supply).

3． Make sure the instrument's power switch is turned off.

4． Connect the AC power cord to the AC LINE end on the rear panel.

5． Use the included AC power cord, or an AC power cord selected by a qualified professional.

6. Plug into an AC power outlet.

1.4 Ground

△ Warning: Make sure the instrument is connected to electrical ground (safety ground, earth).

If the output is connected to the ground that is connected to peripheral equipment or a nearby commercial wire output, and is not directly connected to the ground,

The case may carry very high voltages and become very dangerous.

This instrument is a Class II device (the device is protected against electric shock in addition to basic insulation). However, if not properly grounded

There is still the possibility of electric shock.

To ensure safety, be sure to ground the instrument.

Choose one of the two available methods to ground:

1． The power cord is connected to a single-phase three-wire power socket. (Please make sure the ground wire of the socket is

By connecting the earth)

2． Pass the protective terminal of the rear panel through the grounding bar (the reliable

Copper wire or copper bar) to the ground. Let specialized engineers select, make, and install the interface

Ground connection. To ensure that the ground connection is correct and reliable.

2

1.5 Operation check

△ Warning: When using the instrument normally, make the function interlock (INTERLOCK) as much as possible to ensure the safety of use.

Where the working space is cramped: make a box-like structure for the part under test;

In the case of testing large-scale DUTs with complex structures: use a fence around the test area to protect against electric shock.

When the electric shock protection structure is opened, the (INTERLOCK) signal loop is disconnected to ensure the safety of the workplace.

When the power switch is turned on, the instrument lights up all the lights on the front panel and starts the self-test. Make sure that all the lights are on to ensure safety.

Testing with a DANGER (high voltage danger during test) lamp is particularly dangerous.

△ Caution: After turning off the power switch, wait a few seconds before turning it on again. Repeated on / off power to the instrument without sufficient time interval

damage.

Check order

1． Make sure that the allowable range of the supply voltage is consistent with the input voltage range set by the fuse box.

2． Make sure the AC power cord is connected to the AC LINE end on the rear panel.

3． Insert the power plug into an AC power outlet.

4． Turn on the power switch, make sure that the indicators on the front panel are all on, and the panel displays the startup screen.

5． The next screen displays the AC withstand voltage test (AC) parameter interface of the SETUP interface.

6. Turn off the power switch.

Note the startup screen (example):

The startup screen model, version, and date display are subject to the actual model display.

1.6 Other characteristics of the instrument

(1) Power consumption: power consumption <500VA (RK9920 / A / B) <400VA (RK9910 / A / B).

(2) Overall dimensions (W * H * D): 280mm * 88mm * 420mm;

(3) Weight: about 15kg (RK9920 / A / B); about 13kg (RK9910 / A / B).

| Input voltage | Frequency Range | Fuse (slow melting) | model | rated power |
| --- | --- | --- | --- | --- |
| 110V | 47-63Hz | 5A | RK9920 | 400VA |
|  |  |  | RK9910 | 300VA |
| 220V |  | 3A | RK9920 | 400VA |
|  |  |  | RK9910 | 300VA |

3

Chapter 2 Operating Specifications and Measures

△ Warning: This instrument generates a test voltage of 5 kV that can cause personal injury or even death. When operating the instrument, you must be very careful and observe

2.1 Prohibited operations

■ Do not switch power continuously

After turning off the power switch, make sure to wait one minute or longer before turning on the power switch again. Ensure that the circuit is powered down and restarted.

If the power of the instrument is turned on and off frequently, the control circuit of the instrument may be abnormal due to incomplete power failure. The protection facility may

The protection function cannot be fully implemented.

Except in special or emergency situations. Do not turn off the power switch while the instrument is generating the test voltage.

■ Do not short the output to ground

If the instrument's high-voltage test lead is connected to the AC LINE (AC power line); or other nearby equipment (such as transmission equipment) is connected

The conductor of the earth. When the ground terminal of the instrument is unreliable, the high-voltage low-end (that is, the case of the instrument) will be charged with danger after the high-voltage end is grounded

high pressure.

Make sure the instrument's protective ground terminal and ground wire are connected. This is done even if the HIGH VOLTAGE terminal and the ground terminal are shorted.

There is no danger of being charged with high voltage.

When grounding the protective ground, make sure it is correct and reliable. See "1.5 Grounding".

△ Note: The term "AC LINE" here refers to the power cord used by the instrument. Power for commercial AC or power generation and power for instruments

Connected wires.

■ Do not connect external voltage to the test terminal

Do not connect any external voltage to the output of the instrument. In the non-discharge state, the instrument does not have the function of external discharge. The output terminal and external voltage

Connecting may damage the instrument.

2.2 Handling of emergencies

In the event of an emergency (such as electric shock and burning of the device under test), the instrument does not turn off the high-voltage output. You can do it first

To (a) or (b), but both operations must be done.

(A) Turn off the power switch of the instrument;

(B) Unplug the power cord of the instrument from the power cord plug.

2.3 Preventive measures during testing

■ Wear insulated gloves

Wear insulated gloves when using the instrument to protect yourself from high voltage. Even with high voltage gloves, do not touch live conductors with your hands during high voltage testing.                            4

■ Suspend (suspend) test precautions

When you need to touch the test conductor or change the test connection, press the STOP switch once to ensure that the instrument exits the test preparation state.

If you need to take a break or leave the test place, please turn off the power switch to prevent accidental contact with the start switch.

All dangerous.

■ Charged items during high voltage test

During the test, the high-voltage output terminals, high-voltage test leads, high-voltage probes, DUTs, and exposed conductors around them all carry dangerous high-voltage electricity.

Do not approach or touch these conductors arbitrarily even if there are reliable insulation measures during the test.

△ Warning: The sheath on the alligator clip of the test lead provided by the instrument does not have sufficient insulation for testing high voltage. Do not touch these parts during the test.

■ Precautions after turning off the high voltage output

If you have to touch the device under test, test leads, probes or outputs and the surrounding area due to reconnection or other reasons, make sure

The following two:

(A)

(B)

Confirm that the working status displayed by the instrument is not the test status.

The HV lamp goes out.

■ Remote control warning

Because the start and stop of the high voltage is controlled remotely, the operator cannot know the actual working status of the instrument through the interface, and is performing remote control.

Use caution when using the instrument in mode operation. Pay particular attention to checking the remote control for reliable connections.

(C)

(D)

The "STOP" button must be connected reliably. Before changing the DUT, press the "STOP" button.

When working in a crowded work environment, the remote control switch must have an "INTLOCK" interlock switch and a high-voltage indicator. Replaced by

Disconnect the "INTLOCK" interlock switch before the test piece.

2.4 High voltage test warning

△ Warning: In the high-voltage test, the test leads, test probes, and DUT are charged with high voltage. The instrument has a discharge circuit, sometimes the output is

Discharge is still required after cutting. There is still a danger of electric shock during discharge. To avoid electric shock, make sure the DUT, test leads, probes, and

The output with high voltage does not touch anything other than the test element. If you may touch these, make sure the DANGER light is off, move

In addition to hidden dangers.                    5

Once the test is over, the instrument's discharge circuit begins to force discharge. Do not remove the DUT during the test and before the discharge is completed.

In general, it can be guaranteed that the test loop voltage will be within the safe voltage range when the discharge is completed. When the capacitance of the device under test is too large or

When the component structure is special and will cause incomplete discharge, the technician must change the test method to ensure that the discharge is complete.

■ Discharge time:

Calculation formula of discharge time: t = -ln (30 / U) × R × C

t: discharge time

30: Residual discharge safety voltage 30V

U: Test set voltage

R: discharge impedance of the DUT, the discharge impedance of the instrument is about 10k

C: capacitance of the device under test

Generally, only DC-type high voltage tests need to be discharged. The length of the discharge time depends on the nature of the device under test.

During the test, if it is completed normally, the voltage will drop to zero according to the voltage drop time. If the test fails, the DUT discharge is

It is realized through the secondary side of the transformer (about 10k resistor), and the 1uF capacitor with 6000V high voltage discharges to 30V in about 0.05S. The instrument's fixed discharge time is 0.2S to ensure that the device is discharged.

2.5 Dangerous state treatment of faulty instruments

Typical possible dangerous states of the instrument, the most dangerous of which is “high voltage is output and the instrument is out of control”. When this happens,

1. Turn off the power switch immediately and unplug the AC power cord from the AC power socket.

2. Please stay away from this instrument immediately, and ask the relevant technical personnel to confirm that there is no danger in the test circuit test;

Confirm that there is no output voltage at the test terminal.

3. Remove the relevant connection cables and send the instrument back to us for repair.

△ Warning: Stay away from this instrument immediately after turning off the power, and prevent other people from approaching it. Never disassemble the test circuit immediately.

Contact our distributors or agents immediately. There may be high voltage inside the instrument, problems with non-professionals trying to repair the instrument

Is very dangerous.

2.6 Conditions to guarantee long-term trouble-free use

Due to the volume, weight, and actual use of the instrument, the heat dissipation design of the voltage generation module of the instrument is too small. Therefore, the instrument recommends the following

Use within columns.

Necessary conditions for withstand voltage test

6

| Ambient temperature | Output power | Pause time | Output time limit |
| --- | --- | --- | --- |
| AC T≤40℃ DC | > 12mA (RK9920 / A/B) > 6mA (RK9910 / A/B) | At least as long as the output time | Up to 1 minute |
|  | <8mA (RK9920 / A / B) <4mA (RK9910 / A / B) | No requirement | Can output continuously |
|  | > 6mA (RK9920 / A) > 3mA (RK9910 / A) | At least as long as the output time | Up to 1 minute |
|  | <4mA (RK9920 / A) <2mA (RK9910 / A) | At least one charge waiting time Sample length (WAIT TIME) | Can output continuously |

Note: Output time == (voltage rise time + test time + voltage fall time)

The fan works continuously for 30 minutes, and the use of the instrument must be suspended, otherwise the power amplifier output module may be burned due to overheating.

2.7 Daily inspection

In order to avoid accidents, at least the following points must be guaranteed before use:

1. The input power of the instrument complies with the specifications, and the power supply of the instrument is configured correctly.

2. The instrument is reliably connected to the ground.

3. The test line material is intact without cracks, cracks and breaks.

4. The instrument is not connected to the test lead, and the test is started under the default conditions, and the test can be successfully completed.

5. When connecting the test lead to start the test, the low end of the test lead is in contact with the high end of the test lead, and the instrument can generate a FAIL signal.

Chapter 3 Instrument Panel Overview

This chapter describes the basic operating characteristics of the RK9910 / 20 series instruments. Please read this chapter carefully before using the series instruments

Content so that you can quickly learn the operation of the series.

3.1 Front panel description

Figure 3-1 provides a brief description of the series front panel.

7

Figure Figure 3-1 Front panel description

3.1.1 Off Power switch ((POWER))

switch. Before turning on the power for the first time, the operator should check whether the power type of the instrument and the connection of the test leads are normal.

3.1.2 START key, P STOP key

START key (green circle): used to start the test. Once the test starts, the HV indicator lights up.

STOP key (red circle): Stop key is used to stop the test; it can also be used to cancel PASS, FAIL and other prompt states.

3.1.3 trademark and model

Instrument brand and model

3.1.4

3.1.5 Domain Indicator Area

● FAIL

During the test, if the test data exceeds the set value, the instrument judges that the test is unsatisfactory, and the FAIL judgment light is on.

● PASS

After the test, no test data exceeding the initial setting was found. The instrument judged that the test was qualified, and the PASS judgment light was on.

When the test timing function is turned off (TIME OFF), the test can only be ended with 'STOP' without PASS judgment.

3.1.6 FUNCTION)

Select test mode settings, system settings, file operation interface.

● TEST

When you press this key, the lamp lights, and the instrument enters the test preparation state.

● SETUP

Press this key, the lamp will light, and the instrument will enter the parameter setting interface;

8

● SYSTEM

The button is lit, and the system setting interface (SYSTEM) is displayed

● FILE

The button is lit, and the file operation interface (FILE) is displayed

3.17 HV

DANGER! !! This light will illuminate as long as the test is in progress, indicating that the test is in progress.

3.1.8 High-voltage side of output voltage

High-voltage output of the high-voltage test interface.

3.19 Test low end, test current return end (LOW, RET)

Test voltage output terminal and current sampling terminal.

3.1.10 Move key

It is used to move the cursor on the screen and select parameter items.

3.1.11 Rotary Coded Potentiometer

Used to adjust and confirm parameters.

3.1.12 Shortcut Function Keys

F1-F5 corresponds to the functional operation area on the right side of the LCD for quick operation.

3.1.13 480 * 272TFT dot matrix LCD screen, display setting interface, measurement interface, etc.

3.1.14 USB interface

Implement USB HOST, you can connect peripheral storage.

3.1.15 LOCK

Press this key to lock all keys on the operation panel, except the STARTT and STOP keys. Prevent misuse

3.2 Rear panel description

Figure 3-2 briefly describes the rear panel of RK9910 / 20 series.

Figure 3-2 Rear panel description

9

3.2.1 Amplifier Fan Cooling Port

The heat sink of the power amplifier circuit, please pay attention to keep the space for air circulation.

3.2.2 Test low end, test current return end (option)

The test low end of the spare high-voltage test interface can be modified when the customer needs it.

3.2.3 High-voltage output (option)

The high-voltage output terminal of the spare high-voltage test interface can be modified when the customer needs it.

3.2.4 Power socket

For inputting AC power, please use the voltage within the specified input voltage range of the instrument, please use the power cord that comes with the instrument.

Built-in power fuse, select the fuse corresponding to the input power.

3.2.5 Protective ground terminal

When the three-pin power socket of the instrument power plug cannot guarantee a reliable connection to the earth, it must be connected to a reliable grounding bar from here.

Note: Do not use the instrument without connecting the ground, otherwise the instrument case may be charged with high voltage and there is a danger of electric shock.

3.2.6 Nameplate

Instrument factory serial number record.

Serial communication interface for communication with computer.

3.2.8 USB Device

The U disk mapping of the instrument on the computer side can be realized, and the internal storage files can be directly read.

3.2.9 HANDLER interface

24V power supply, INTLOCK function.

TEST: The synchronous control signal output by the machine when the instrument starts the high voltage output.

START: Input the start signal of this unit to start the high voltage output, which is equivalent to the START signal on the front panel.

RESET: Input the reset signal of this unit to stop the high voltage output, which is equivalent to the STOP signal on the front panel.

PASS: The qualified signal output by this machine is equivalent to the PASS indicator on the front panel.

FAIL: The failure signal output by this machine is equivalent to the FAIL indicator on the front panel.

3.2.10 SINGLE interface

This interface is an online protection and internal 24V power output interface.

INT LOCK:

(Port: (5) INT LOCK + (6) COM) Short circuit is valid.

This machine's online lock signal is shorted by short-circuit jumper by default. The unit is not allowed to start the test when disconnected.

10

DC 24V power supply:

(Ports: (1, 2) 24V + (3, 4) GND).

The output voltage is 18.5VAC rectified output, without voltage regulation function. It is used for the control power of indicators, etc. It is recommended that users use

Current <500mA.

3.2.11    110 / 220V power transfer switch

Realize the conversion of input power voltage 110V 220V

3.3 Instrument Performance Overview

RK9920 can provide 5kVAC / 20mA withstand voltage, 6kVDC / 10mA withstand voltage, insulation resistance test.

RK9920A can provide 5kVAC / 20mA withstand voltage and 6kVDC / 10mA withstand voltage test.

RK9920B can provide 5kVAC / 20mA withstand voltage test.

RK9910 can provide 5kVAC / 10mA withstand voltage, 6kVDC / 5mA withstand voltage, insulation resistance test.

RK9910A can provide 5kVAC / 10mA withstand voltage and 6kVDC / 5mA withstand voltage test.

RK9910B can provide 5kVAC / 10mA withstand voltage test.

Principle structure of the instrument: Since the MCU processor generates a signal, a sine wave is generated by the signal processing circuit and amplified by a power amplifier.

High voltage is produced by 40 ～ 600Hz high voltage transformer. The DA reference signal guarantees controllable voltage amplitude. Controllable sine wave can be set during AC output

Set at 50 or 60Hz. Aiming at the problem of power supply ripple in DC and insulation resistance tests, a 600Hz AC power rectifier is used to form a DC.

Voltage as power. The output voltage is closed-loop controlled to ensure a small load regulation.

RK9910 / 20 series can perform AC withstand voltage test, DC withstand voltage test, and insulation resistance test.

RK9910 / 20 series are equipped with HANDLER, RS-232C, RS485, USB, making the instrument adaptable to many different

Automated test systems with high safety and reliability are required.

Features:

■ Three test functions—AC withstand voltage test, DC withstand voltage test, insulation resistance test,

RK9910, RK9920 provide AC and DC withstand voltage test and insulation resistance test.

RK9910A and RK9920A provide AC and DC withstand voltage test.

RK9910B and RK9920B provide AC withstand voltage test.

■ Two test power options

The RK9920 series high-voltage module is a class AB power amplifier circuit and a 100VA high-voltage transformer, achieving AC, 5kV / 20mA

Output; DC, 6kV / 10mA output. The distortion of the waveform is less than 3%.

11

The RK9910 series high-voltage module is a class AB power amplifier circuit and a 50VA high-voltage transformer, achieving AC, 5kV / 10mA

Output; and DC, 6kV / 5mA output. The distortion of the waveform is less than 3%.

If the customer is continuous current output, in order to ensure the reliability of the instrument, when the maximum output is working when it is more than 60% of the rated output current

Time is 60 seconds. Within 60% to 40% of the rated output current, please pay attention to limit the continuous working time. Available below 40% of rated output current

To ensure continuous work.

Figure 3-3 AC voltage load regulation

■ DC withstand voltage test 6kV / 10mA (RK9920 / A) 6kV / 5mA (RK9910 / A)

RK9910 / 20 series can provide DC withstand voltage test with wide voltage range (maximum output DC 6kV). 600Hz frequency hardware

Automatic voltage adjustment, voltage load adjustment rate ≤1% + 10V.

5.00

4.00

3.00

2.00

1.00

0          2        4         6       8         10

V/mA

Figure 3-4 RK9920 / A  DC voltage output range

■ Insulation resistance test 0.050kV to 1.000kV (1V resolution) /0.1MΩ to 10.0GΩ, maximum rated current RK9920

10mA, RK9910 is 5mA.

Insulation resistance test range:

When the voltage is less than 500V: 0.1MΩ ~ 1GΩ The accuracy is ± [10% of reading + 5 words]

When the voltage is greater than 500V: 0.1MΩ ~ 100MΩ range accuracy is ± (5% of reading + 5 words), 100MΩ ~ 1GΩ range precision

The degree is ± (10% of reading + 5 words).

12

■ RS-232C interface as standard

In addition to power conversion, key lock and other functions, all other can be controlled remotely. In DC withstand voltage test, AC withstand voltage test,

In the insulation resistance test, test conditions such as test voltage, judgment function, and test time can be remotely controlled. Test results can also be passed remotely

The control is read from the back. RS-232C, RS485 interface provides stable and unified standard test with PC or other equipment

■ HANDLER interface and input signal interface for convenient connection and control

HANDLER interface: You can input START, STOP signals, and output TEST, PASS, FAIL signals. Can be very convenient

The foot switch is connected to a foot control, which is connected with a simple test fixture to achieve safety interlock, pneumatic control, test instructions and so on.

Input signal interface: It can input INT LOCK signal and provide 24V, 0.5A power output at the same time, which is convenient for control and connection.

■ USB interface for backup

The instrument is equipped with a USB interface, which can save the test plan and customer measurement files written by the instrument to an external USB flash drive, or

Load the instrument to facilitate the batch setting of the instrument's operating parameters and archiving scheme.

■ Test waiting time setting

The test waiting time can be set from 0.1s to 999.9s with a resolution of 0.1s. During this time, the instrument will output a TEST control signal to control

Control external equipment to ensure the test connection is reliable, and then start the high-voltage test process.

■ Rise time control function

During AC withstand voltage test, DC withstand voltage test and insulation resistance test, the test voltage can slowly rise to the set value instead of

Immediately after starting the test, the set voltage is supplied to the device under test. The voltage rise time is 0.1s to 999.9s with a resolution of 0.1s. TH9310 / 20

The series complies with various UL test standards and IEC withstand voltage test standards (the initial voltage is less than half of the test voltage and

The test voltage can be specified for the rise time).

■ Drop time control function

In the qualified judgment in the AC withstand voltage test, the test voltage can be gradually reduced. Voltage drop time can be set from 0.1s to 999.9s

Resolution between 0.1s.

■ Discharge function

The DUT is usually capacitive. When the DC withstand voltage test and insulation resistance test are cut off, the DUT remains fully charged.

There is a danger of electric shock. The RK9910 / 20 series has a strong resistance to the DUT after the DC withstand voltage test and the insulation resistance test are completed.

Control fast discharge function.

■ Enhanced Security

13

In order to improve safety, RK9910 / 20 series is equipped with many facilities and safety

functions, including safety output, discharge function and ground current detection.

The so-called ground current detection is to cut off the high voltage output when the return current of the local high voltage test circuit through the case is greater than 0.45mA.

■ High test accuracy

RK9910 / 20 series voltage digital display. The voltage test accuracy is ± (1% of reading + 5V) during the withstand voltage test.

Voltage accuracy is ± (1% of reading + 2V). The accuracy during the withstand voltage and current test is ± (1% of reading + 5 words).

■ Simple operability

The RK9910 / 20 series is easy to operate, ensuring that users do not have difficulty getting started. This instrument lists all tests in the setting interface

parameter. Use the arrow keys to select a parameter from the interface displayed on the LCD, and then use the function button to modify the parameter.

To take measurements directly.

■ 16M FLASH, each can have 50 test items.

20 test files can be edited, corresponding to various test project combinations required by customers; each test file can have up to 20 test items

The test items are AC withstand voltage test, DC withstand voltage test, insulation resistance test, and the test conditions of each item are irrelevant.

The stored files of the instrument can be transferred to a computer or another computer via an external USB flash drive in the file operation interface.

Option function description:

■ High-voltage output on the rear panel (optional)

The rear panel includes an optional high-voltage output. This port provides a convenient solution for wiring instruments in a cabinet.

!! Warning: This instrument uses a high voltage of 5kV AC / DC. Therefore, do not touch the DUT and the test lead at random, it may cause an electric shock.

Around the device under test, safety measures similar to fences should be provided to ensure user safety. In addition, to ensure safety,

Extreme care is taken to prevent improper connections and high output voltages.

14

Chapter 4 Basic Operations

4.1 Instrument Interface Structure Overview

This chapter describes the operating procedures for the withstand voltage and insulation resistance tests. The interface structure of the instrument is shown as follows:

TEST

Start test

File modification and inspection

DC（DC withstand voltage）

IR（Insulation resistance）

STEP

SETUP

SYSTEM1

FILE

AC（AC withstand voltage）

AC  parameter

SYSTEM1

SYSTEM2

SAVE

LOAD

Test operation flow diagram

Interface description:

4.1.1 The first column of the interface structure is written based on the initial state of the panel function keys.

Detailed description). TEST interface cannot modify parameters.

4.1.2 The first interface structure The second column is the parameter structure of the initial interface. For example, P SETUP interface default STEP: program step 1 1, total step number 1 1,

AC: AC withstand voltage test interface, C AC parameters: other parameters are AC withstand voltage. Test parameters.

4.1.3 Interface structure The third column is the function switching interface. In the second interface, some functions can be changed when some function marks are selected. This interface

The related parameters will change. If C AC is changed to DC, the instrument will change the AC withstand voltage test mode to DC withstand voltage test mode.

The 'C AC parameter' on the previous interface will become the 'C DC parameter' which needs to be set for the DC withstand voltage.

1 1. Press and hold the STOP and START keys simultaneously to turn on the power again, and the instrument will restore the factory default settings.

2 2. The clear range is: (SETUP) test condition, (SYSTEM) system setting.

3 3, When the instrument software is upgraded, or when the display error is caused by calling a low version file, you can use this method to restore the instrument

Device works normally.

15

4.2 Panel Function Interface and parameter description

This section mainly introduces the function interface and related parameters of the instrument based on the software process and interface correlation. To guide users through the capabilities of the instrument

Can and use.

Instrument initial state description

1. After booting, the system defaults to: the last modified setting interface.

2. The factory setting of the instrument is single step, AC withstand voltage, default parameter state. As shown below .

3. The default cursor of the default interface is interface switching.

Instrument default interface diagram

The interface of the instrument can be switched directly with four function keys, which are measurement display (TEST), measurement setup (SETUP),

System interface (SYSTEM), file processing (FILE) The following describes the functions of the interface.

Basic function description of panel keys:

TEST (test key): Puts the instrument into the test waiting state and prepares to start the high voltage test.

SETUP (setting key): Interface to modify the current test plan, test items, and test parameters. Changes to the test plan

It's all done in this interface.

SYSTEM (system key): The settings related to test safety and instrument operating mode have little to do with the test plan.

FILE (file key): Save and recall interface of test plan, related to data storage.

(Arrow key): The cursor can be moved freely between parameters.

F1 to F5: (Soft key): Cooperate with the contents of the soft key function display area to modify the object selected by the cursor.

Instrument system upgrade steps:

16

4.2.1 P SETUP measurement setup. The interface is as follows:

Figure 4.2.1 AC setting interface

Test plan modification

STEP: 01/01 Test steps: current set item number / total

The project ID of the test plan. The current parameter is the number of items of the test plan.

| Button function | Chinese | description |
| --- | --- | --- |
| F1 NEW | New | Create a new empty test plan (STEP), the system will automatically create a new default test project. Pay attention to saving after writing your own test plan. |
| F2 NEW | insert | A new test project was added after this project. The current and subsequent items will move backward One person. |
| F3 DEL | delete | Delete the current test project. Subsequent items will be moved forward by one. |
| F5 Save as file |  | Save test parameters |

The working mode of the current test project is AC withstand voltage. When the cursor is in this position, it can be switched to DC, ， IR, ， project through the coded rotary potentiometer.

For details of parameter modification of the test item, please refer to “4 4. .3 3 Test Item Interface and Parameter Description”.

note:

For ease of use, when "Test Protection" F OFF.

Press ‘START’ key in this interface to directly enter the test interface, and start the test according to the currently set test scheme.

If the operation of “START” may be triggered by mistake, please unplug “INTLOCK” before modifying the parameters.

17

4.2.2  TEST test interface. The interface is shown as follows: (taking  AC as an example)

Figure 4.2.2 AC test interface

Note: 1. In this interface, you can start high voltage to perform high voltage measurement on the component under test. Its test parameters must be set in detail in the setting interface.

2. Press F2 in this interface to quickly switch to the SETUP interface.

3. The LOCK function key can lock the keyboard. After the keyboard is locked, the instrument only responds to the START, STOP, and (unlock) keys.

After starting the measurement, the middle of the instrument panel displays three data in large font. Display the real-time test data during the test.

The result of the last test is displayed before the STOP key.

The upper one is the high-voltage output voltage and the withstand voltage is in kilovolts (kV).

The middle one is the low-end measured current, measured in milliamps (mA) and microamps (uA).

The lower one is the remaining time of the corresponding step during the high voltage test. If the user turns off the test time, the test time shows that the test is entered

The time after the state. This count is no longer accumulated after the count is greater than 999.9. If it is not "FAIL", the test status must be exited with "STOP". Users can analyze the test situation of the tested objects intuitively. In seconds (S).

In general, if “FAIL” occurs, the instrument will wait for “STOP” to exit the L FAIL state.

For ease of use, when "Test Protection" F OFF. The instrument will automatically exit the "FAIL" state. Please pay special attention when using it.

!! special reminder:

When the instrument is switched on, the operator must not leave the test station! !! !!

During the high-voltage test, no one can access the test line of the instrument or the device under test! !! !!

Non-operator personnel must not approach the withstand voltage instrument that is turned on to avoid danger! !! !!

18

4.2.3 SYST E M system interface. The interface diagram is as follows:

The system interface is to set some settings that have nothing to do with the specific test component parameters, but are related to the test plan of the instrument.

1.SYSTEM1 interface

<System Settings 1>

Interface parameter description:

| english | Set flag | Parameter value description | meaning |
| --- | --- | --- | --- |
| Failure mode | FAIL MODE | STOP | Exit directly when not qualified |
|  |  | CONTINUE | Continue the test when it fails, and the results are output on the form |
|  |  | REST | If it fails, press START to retest Current failed step |
|  |  | NEXT | If it fails, press START to continue testing step. |
| Electric shock protection | GFI： | ON 、OFF | Ground current detection, enable |
| Item-to-item hold | STEP HOLD | 0.2S～99.9S | Wait time between projects during multi-project testing. |
|  |  | OFF | Don't wait between projects |
|  |  | KEY | Pause, press ‘START’ to start the next item |
| Test protection | TEST SAFE | ON | FAIL status must be returned to wait status with STOP, In order to start the next test with START. |
|  | OFF | FAIL status automatically returns to wait status, use START Start the next test directly. |
| Control mode: | CTRL MODE | FILE | When the test file ends, the HANDLER interface outputs the test. Test results. |
|  | STEP | At the end of each step, the HANDLER interface outputs Test results from previous steps. |
| Qualified voice | PASS BEEP | OFF | shut down |
|  | ON | OPEN |

19

| Key sound | KEY BEEP | OFF | Turn off the button |
| --- | --- | --- | --- |
|  |  | ON | On button sound |
| Zero setting | OFFSET | OFF | Start testing as soon as the default instrument is ready |
|  |  | OFF~ON | The base number is set to zero. |
|  |  | GET | Get the test base under the current test conditions. |
| Short circuit threshold | SHORT | LOW | Short-circuit response sensitivity: low |
|  |  | HIGH | Short-circuit response sensitivity: high |
| Bus address: | BUS ADDR | 0-31 | Address when the instrument is transferred to the bus |

<System Settings 2>

SYSTEM2 interface

Interface parameter description:

English      Set flag      Parameter value description         meaning

Bus mode    BUS MODE        RS232C           Serial mode: Data format: 8.n. 1

RS485

Baud rate     BAUD             9600~115200       Serial bus baud rate.

default setting  DEFAULT                       Restore the factory system settings.

LANGUAGE    LANGUAGE      chinaese/English   Instrument interface language

selection.

Password setting   PASSWORD     OFF       Disable the password function. The default password is 9920.

SYSTEM   A password is required to boot into the system.

FILE      File calls require a password.

Modify    Change the old password to the new password.

20

4.2.4 FILE file storage interface.

Press (FILE key) to enter the file management interface as follows:

FILE interface

Interface structure description:

NO.       Description        Quick options          meaning

1        Internal storage                      Internal file interface

Extenal storage                      External file interface

2        document list             F1    Call the current file as a file for internal use

F2    Copy the current file to a USB drive

F3    Delete the current file

F5    Refresh display

File list page up

File list page down

21

4.3 Test project interface and parameters

This section introduces the test function parameters of the setting interface and their

meanings to guide customers in setting related parameters.

4.3.1  AC withstand voltage test parameter setting. The setting interface is as follows:

Figure 4.3.1 AC setting interface

Description of AC withstand voltage (AC) test parameters:

| VOLT ： | 0.050～5.000kV | AC high voltage test voltage |
| --- | --- | --- |
| UPPER ： | 0.001～20.00mA | RK9920 series AC withstand voltage and current upper limit |
|  | 0.001～10.00mA | RK9910 series AC withstand voltage and current upper limit |
| LOWR ： | 0.001～20.00mA | R K9920 series AC lower limit current value, Must be less than the UPPER value. |
|  | 0.001～10.00mA | R K9910 series AC lower limit current value, Must be less than the UPPER value. |
|  | OFF | No lower limit required |
| ARC ： | 1.0 ～ 20.0 mA | Maximum allowable AC arc current |
|  | OFF | No arc required |
| TIME ： | 0.1 999.9S | AC withstand voltage test time, the test ends when the time is up RISE ≠ OFF |
|  | OFF | Unlimited test time |
| RISE ： | 0.1～999.9S | AC high voltage test voltage rise time |
|  | OFF | Default = 0.1S, test time> 0.2S. |
| FALL： | 0.1～999.9S | AC high voltage test voltage drop time |
|  | OFF | Directly cut off the voltage output after the test. (DUT may be charged) |
| FREQ ： | 50/60 Hz | AC working frequency |

22

4.3.2 DC withstand voltage test parameter setting. The setting interface is as follows:

Figure 4.3.2 DC setting interface

The DC withstand voltage (DC) test parameters are explained as follows

| VOLT ： | 0.050～6.000kV | DC high voltage test voltage |
| --- | --- | --- |
| UPPER ： | 0.1 uA～10.00mA | RK9920 / A DC withstand voltage and current upper limit |
|  | 0.1 uA～5.00mA | RK9910 / A DC withstand voltage and current upper limit |
| LOWR ： | 0.1 uA～10.00mA | RK9920 / A DC withstand voltage and current lower limit, Less than UPPER value. |
|  | 0. 1 uA～5.00mA | RK9910 / A DC withstand voltage and current lower limit, Less than UPPER value. |
|  | OFF | No lower limit required |
| TIME： | 0.1～999.9S | DC withstand voltage test time, the test ends when the time is up, RISE ≠ OFF |
|  | OFF | Unlimited test time |
| RISE： | 0.1 999.9S | DC high voltage test voltage rise time |
|  | OFF | Default = 0.1S, test time> 0.2s |
| FALL： | 0.1- 999.9S | DC high voltage test voltage fall time |
|  | OFF | At the end of the test, the voltage output is directly cut off, and it enters 0.2S to discharge quickly. |
| WATT： | 0.1～999.9S | DC charging wait time. |
|  | OFF | Components do not need to be charged |
| ARC: | 1.0 ～ 20.0 mA | DC Arc Current Max |
|  | OFF | No arc required |
| RAMP ： | ON | Voltage rise time and current upper limit judgment are allowed. |
|  | OFF | Voltage rise time, does not detect upper current limit, but The current limit judgment is still judged. |

23

4.3.3 IR insulation resistance test parameter setting. The setting interface is as follows:

(Schematic 4.3.3):

IR setting interface

The insulation resistance (IR) test parameters are explained as follows:

| VOLT ： | 0.050～1.000kV | Insulation test voltage value. |
| --- | --- | --- |
| UPPER ： | 0.1 M～10.00G | Upper limit of insulation resistance. |
|  | OFF | Do not judge the upper limit of insulation resistance |
| LOWR ： | 0.1M～1 0.0G | Lower limit of insulation resistance, less than UPPER value. |
| TIME： | 0.1～999.9S | Insulation resistance test time. (RISE ≠ OFF) |
|  | OFF | The test time is unlimited. |
| RISE： | 0.1 999.9S | Rise time of insulation voltage. |
|  | OFF | Default = 0.1S, test time> 0.2s. |
| FALL： | 0.1- 999.9S | Insulation voltage fall time. |
|  | OFF | After the test, directly cut off the voltage output and enter 0.2S fast discharge. |
| RANG： | AUTO | Auto range mode: improve test accuracy. |
| RK9920 series | 1.0 ～ 20.0 mA | DC Arc Current Max |
|  | OFF | No arc required |
| RK9910 series | 2uA、20uA 、200uA 、2mA 10mA |  |
|  | 1uA、10uA 、100uA 、1mA、5mA |  |

note:

1,Only SHORT FAIL will respond during charging and only testing. At the end of the test,

the lower test of the test will be judged. Please note! !! !!

2, RANG AUTO, due to range switching, the minimum test time is 0.6S, please note! !! !!

Because the lag of current acquisition is about 20mS compared to voltage acquisition,

3,the voltage rise time shows that the resistance value is too small, and the voltage fall time shows that the resistance value is too large, hereby explain for reference.

24

4.4 Test function principle and instructions

This section introduces the principles and uses of tests such as ground connection, ground current detection, and arc detection in the order of the test process.

Test flow diagram of the instrument

1,Start  test

2,Test delay

3,Voltage rise

4,Charging current detection

5,high-pressure test

6,Voltage drop

10,Test  result  processing

12, End of test

Does

not

handle

current

overruns

The fllowing

Unqualified

Judgments

Will

Be

Triggered

If

unqualified

7,GFI

ARC

SHORT

8,HI,LOW

9,Disqualification

11,STOP

4.4.1 Start test

After the instrument is in the test mode, check the test conditions and the DUT is connected correctly, and then press the START key to start the test.

4.4.2 Delay Test Delay

25

4.4.3 Voltage rise

Some DUT characteristics are more sensitive to sudden changes in voltage and need to use this function. When the instrument starts to output, the output voltage is zero. When the instrument starts to output voltage, the instrument will control the output voltage to rise in steps of 0.1S. The step-up value is based on the test voltage and voltage rise time.

OK (ΔV = V / (10 * S)). If you turn off the voltage rise time (RISE OFF), the default voltage rise time of 0.1 seconds is automatically added to the test time so that the minimum test time is 0.2S. If this value is too small, it may cause ARC or DC boost determination error. Please note.

4.4.4 DC boost determination

Whether to activate the current upper limit judgment function switch during the voltage rise is mainly used to avoid test misjudgment. When the test distributed capacitance is small and the charging current of the capacitor is relatively small, which will not cause a significant change in the current, turning on the boost judgment can detect the poor performance of the DUT early, reducing the chance of component overcurrent damage.

When the step capacitance is large, the capacitor will have a charging process during the voltage rise. At this time, the current may be much larger than the upper limit of the measured current. If the boost determination is turned on, it will cause a misjudgment of the upper limit. If necessary, you can open the short circuit threshold to lower the short circuit sensitivity and increase the charging current.

4.4.5 High voltage test

Perform a high voltage test on the device under test. At this time, it should be possible to ensure that the test circuit is correct, the test result will not be affected by some special incidental parameters, and the display content is the actual withstand voltage and current required for the test.

4.4.6 Test voltage drop

The rise in test voltage is determined by the characteristics of the DUT. When the voltage drops at the end of the high-voltage test, the instrument will control the output voltage to decrease in units of 0.1S (the DC voltage will not decrease with the control voltage). * S)). If you turn off the voltage drop time (FAIL OFF), the default voltage drop time is 0.1 second. At this time, the instrument does not make test comparison judgments, and the data is for reference only. When the voltage drop is over, the instrument will connect the test circuit to the AC withstand voltage mode. At this time, if the DUT has a DC voltage drop, it will discharge through the instrument's AC circuit.

4.4.7 Ground Current Detection Function

Ground current detection is to detect the current flowing through the instrument case to prevent electric shock. When high voltage is output, there is current flowing back from the voltage output terminal through the human body to the instrument casing, which may cause very serious consequences.

Ground wire current detection and judgment circuit response description of the instrument:

Ground line current detection is enabled. If the ground line current is greater than 0.45mA, it is judged that the ground line current exceeds the limit.

26

When judging the electric shock, the instrument will end the high voltage output within

0.3S and exit the test state. And (GFI FAIL) is displayed.

Note: The instantaneous output current of the instrument may be greater than 30mA. If it is indeed an electric shock, it may cause the operator to comatose or die. Therefore, it is recommended to enable ground current detection when the product allows it.

4.4.8 Current overrun and arc detection

Current over-limit classification: current lower limit, current upper limit, current over-limit, arc detection.

Current lower limit judgment (LOW): Generally used for testing the low-end disconnection judgment. When the instrument tests the equipment, the equipment must have a certain leakage current. When the leakage current of the instrument is less than the lower limit set current value, the test is considered to fail (no equipment is connected).

If the leakage current of the component under test is small, this function must be turned off. Judgment display (LOW FAIL) when the limit is exceeded, only test mode, this judgment is valid, timing sampling, the rate is 100mS each time.

Current upper limit judgment (HIGH): The most commonly used test current exceeds the limit judgment. When the instrument tests the equipment, the equipment must have a certain leakage current. When the leakage current measured by the instrument is greater than the upper limit set current value, it is considered that the equipment's withstand voltage and impedance are not enough to fail the test. Judging and displaying when exceeding limit (HI FAIL), timing sampling, the rate is 100mS each time.

Current limit judgment: The current sampling judgment is slower, the current changes faster when the insulation is broken, and the sampling circuit cannot respond in time, and the current peak exceeds the output range allowed by the instrument. FAIL).

Because the data cannot be collected after such a current exceeds the limit, the result output by the system at this time is: the test result within 100mS before the current exceeds the limit. The current limit is twice the permissible output current of the instrument (relative AC is 1.5 times the peak value). The fall time is invalid, this judgment cannot be masked.

Arc detection (ARC): It is a very practical function for the measurement of coil components. It tests in the high-voltage test circuit.

There is a local current oscillation caused by a partial circuit instantaneous discharge. Because it is superimposed on the normal test current, the mutation time is short.

The ordinary current detection circuit on the board cannot make a proper judgment in response to the current change. The arc detection circuit filters out the normal current value,

Handles only high-speed current pulse changes. Due to the random nature of the low-pass filtering and the size of the arc, this function can only be approximated.

Calculate the extent of local fire. Because the data cannot be collected after the current exceeds the limit, the output result at this time is the last test when it passed.

As a result, the judgment display (ARC FAIL) is displayed. ARC current is a qualitative analysis for testing. The amount of measurement and test environment,

Test line distribution and so on have a great impact on randomness, so pay attention when

27

using it.

1K     10k      100k

Frequency response comparison between current overrun judgment and arc detection: (see the figure above)

Area A in the figure: shows the frequency response of the current sampling circuit. Because it is necessary to filter the ripple of the power frequency → AD sampling → Calculate the test result → Analyze whether the current exceeds the set limit. Within the test current range, the pulse width is greater than 100mS.

Area B in the picture: the current fast response circuit. It only filters out high-frequency interference signals → voltage peak comparison → overcurrent peak signal lock, and only makes limit judgments. It is greater than the allowable output current of the instrument, and the pulse width is greater than 1mS.

Area C in the picture: Arc detection circuit. The arc detection circuit only samples the amplitude of the sudden change in the current, and the signal high-pass filters out the low frequency → voltage peak comparison → pulse lock. The pulse width is about 1uS-1mS at the sudden edge of the current near the set value.

4.4.9 Non-conformance judgment

1. During the test, the current exceeds the maximum output current that the instrument can withstand, or if the instrument finds a state of potential safety hazards, the instrument will immediately cut off the voltage of the test circuit and wait for the instrument software to query the cause of the error to determine the failure.

2. If the test result exceeds the limit set by the test item, the instrument will judge that the test piece is unqualified. Immediately stop the current test, cut off the voltage output, and enter the processing procedure of unsatisfactory test results.

3. Multi-step test. There is a step of FAIL, and the overall test result is FAIL.

4.4.10 Test result processing

If the test process exceeds the limit, it is judged as (FAIL). If FAIL appears in a multi-step test, the final result is FAIL. There are multiple test items, and FAIL's judgment processing mode is controlled by the system's failure mode. Otherwise, the instrument will display FAIL judgment and category (see the figure below for HI as an example), waiting for the user to process.

After the test, there is no disqualification mark, and the test result is judged as (PASS).

The PASS judgment processing mode is controlled by the PASS HOLD of the SYSTEM, and then it is ready to start the next measurement or return test

Waiting.

FILE mode, then the test will be output only after the entire file test is finished

The HANDLER signal output is controlled by the control mode. Select FILE mode, then the test results will be output only after the entire file test is finished. The STEP mode

28

controls the interface to output corresponding signals at each step.

From this state until the next step, customers can use software to query test data and results

4.4.11 STOP (Stop Measurement)

Press the 'STOP' key in any state during the entire test process, the instrument automatically ends the test and enters the test end state. Press the “STOP” key again, the instrument will return to the test waiting state. No test result judgment output will be given when the test is stopped.

At the end of the test, the customer can use software to query the last test data obtained before 'STOP'.

4.4.12 OFFSET

Before testing, due to changes in the working environment of the instrument and the placement of the test cables, some base numbers may appear during the no-load test of the instrument. For customers who require precise measurements, they can be cleared in the SYSTEM interface.

The specific steps are as follows:

1,Set the current test conditions on the SETUP screen.

2． Select the OFFSET item on the SYSTEM interface and set it to ON.

3． Press GET, the instrument will automatically start the high voltage test and use the current test value as the zero value.

4． If the customer has not set the test time, the test can be stopped with the STOP key.

Note:

Do not connect the product to the test terminal during GET. Zeroing the product cannot increase the actual test range of the instrument.

OFFSET will clear all steps, please confirm the setting data is correct.

To change the test conditions, OFFSET must be performed again.

29

4.5 HANDLER and SIGNAL interface circuit structure and use

4.5.1 Principles of Control Interface

The internal principle of the HANDLER and SIGNAL interface instruments is as follows:

HANDLER, SIGNAL interface structure and timing

Explanation:

1, HANDLER interface: START, STOP, COM signals constitute remote input control, switch input closed effective.

2. HANDLER interface: TEST, PASS, FAIL signals form remote output control. The switch output is closed.TEST can be used as a high-voltage start signal or a pulse signal when the instrument is working properly.

3. The SIGNAL interface is mainly used to provide the instrument selection signal (INTLOCK) for multi-instrument online test. This signal is a short-circuit by default. When the circuit is open, the instrument is prohibited from starting high-voltage output.

4. The SIGNAL interface also provides a power source with an output voltage of approximately + 24V and an output current of less than 0.5A. With the HANDLER interface control signal, it can be used to drive indicators, photoelectric switches, low-power solenoid valves, and so on. (See below)

30

4.5.2 Control Interface Use

The control interface is generally used for remote control and test synchronization or indication. The external connections of the interface are as follows:

Figure 4.5.2: Schematic diagram of external circuit connection

Explanation:

1. The switch can be replaced with an isolated switching element such as the optocoupler secondary side. The current direction is based on the principle shown above (the COM terminal is the low side).

2. The indicator light can be replaced with other drive control components. The current direction depends on the power supply.

3.Internal power performance of the instrument

a) The rectified and filtered output of the AC power supply. The unregulated output is about 24V. Please confirm before use.

b) The instantaneous maximum current value should not be greater than 0.5A, and the long-term working current should be less than 0.2A. If you need more current, please bring your own power supply.

c) The external control signal needs to be greater than 220V voltage or 2A current, the internal relay of the instrument will not be able to bear it, please transfer it by yourself.

4.6 Other interfaces and functions of the instrument

1. On the front panel, USB HOST is used to connect the USB flash drive, and it is used for export and import of customer setting files.

2. Rear panel, USB DEV maps internal FLASH to computer as U disk.

3. RS232 is used to connect with the computer. For the baud rate, see the system setting item. The data format is 8.n.1. Compatible software format IEE485.

31

Chapter 5 Description of Serial Instruction Set

Brief description of the instruction format:

1. The instrument instruction set only describes the actual characters that the instrument accepts or sends.

2. The instruction characters are all ASCII characters.

3. The command data "<???>" are all ASCII strings. The default format of the system is integer or floating point number. The unit of data is the default value.

4. The end of the instruction must have the end of instruction mark: an identifier of the end of the instruction, without which the instrument does not parse the instruction.

a) The default end tags are: carriage return (N L), print control (\ n), decimal number (10), hexadecimal number (0x0A).

b) End mark of IEEE-488 bus: keyword (^ END), signal (EOI).

Multiple instructions can simplify sending examples as follows: Note: "_" in the example is a space mark

FUNC：SOUR：STEP_1：AC：VOLT_1000；UPPC_1；TTIM_9.9；CH1_HIGH；CH2_LOW (NL^END)

FUNC：SOUR：STEP_INS (NL^END)

FUNC：SOUR：STEP_2：DC：VOLT_1000；UPPC_1；TTIM_9.9；CH1_HIGH；CH2_LOW (NL^END)

5.1 SCPI instruction set

RK9920 / 9910 instrument subsystem commands

●Display   ●function

●system ●mmen ●FEIC

5.2 DISPLAY subsystem command set

The DISPlay subsystem command set is mainly used to set the display page of the instrument. Character? Query returns the current page.

DISPlay: PAGE

Command syntax: DISPlay: PAGE <page name>

<page name> is as follows:

TEST, set display page to: measurement display page

TESTSET, set the display page to: measurement setting page

SYSSet, set the display page to: system setting page

FLIE, set the display page to: (internal) file list

character? Query returns the current page.

--example:

Set the display page to: Measurement display page.

Setting instruction: DISP: PAGE TEST

Query instruction: DISPlay: PAGE?

Return value: TEST

5.3 FUNCtion subsystem command set

5.3.1 FUNCtion subsystem command set is mainly used to set the test parameters of the instrument test function.

Command tree.

32

FUNCtion    :STEP: ---- num:      :AC	:VOLTage

:UPPC

:LOWC

:STARt                     :TTIM

:RTIM

:STOP                      :FTIM

:ARC

:FREQ

:DC    VOLTAGE

UPPC

LOWC

TTIM

RTIM

FTIM

ARC

RAMP

:IR     VOLTAGE

UPPR

LOWR

TTIM

RTIM

FTIM

RANG

NEW

DEL

5.3.2 PROG Function Command Set

FUNC：STARt, When the instrument is in the test interface, start the test

FUNC：STOP, Stop the test when the instrument is in the test interface

FUNC：STEP :<num>:INS , Add a new test project to the existing test plan (STEP)

FUNC：STEP :<num>:DEL, Delete the current test item in the existing test plan (STEP)

FUNC：STEP :<num>:NEW, Create an empty test plan to write a new test plan

FUNC：STEP :<num>:Edit the <num> step of the current test scenario, <num> = 1 ~ 50

33

5.3.3 AC Setup Function Command Set

FUNC：SOURce：STEP：AC：VOLT       Set / Query ACW voltage

--format:

Setting format: FUNC: STEP: <num>: AC: VOLT <voltage value>

Query format: FUNC: STEP: <num> :: AC: VOLT?

--Data <sn>

Data type: integer

Data range: 1 ~ 20

Data accuracy:1

--Data <voltage value>:

Data type: floating point

data range:

0.050 ~ 5.000

Data accuracy: 0.001

--Example: Data unit: KV

Set the ACW voltage in STEP 1: 1000V

Set command:

FUNC: STEP: <num>: AC: VOLT 1.000

Query command: FUNC: STEP: <num>: AC: VOLT ?,

Return value: 1.000

FUNC: SOURce: STEP: AC: UPPC, Set / Query ACW upper limit current

--format:

Setting format: FUNC: STEP: <num> :: AC: UPPC <current value>

Query format: FUNC: STEP: <num> :: AC: UPPC?

--Data <voltage value>:

Data type: floating point

Data range: 0.001 to 20.000 mA

Data accuracy: 0.001 mA

Data unit: mA

--example:

Set the ACW current limit in STEP 1 to: 1mA

Setting command: FUNC: STEP: <num> :: AC: UPPC 1

Query command: FUNC: STEP: <num> :: AC: UPPC?

Return value: 1.000

FUNC: SOURce: STEP: AC: LOWC

--format:

Set / Query ACW lower limit current

Setting format: FUNC: STEP: <num> :: AC: LOWC <current value>

Query format: FUNC: STEP: <num> :: AC: LOWC?

--Data <current value>

Data type: floating point

Data range: 0 ~ 20.000 mA (where 0 is OFF) RK9920

0 ~ 10.000 mA (where 0 is OFF) RK9910

Data accuracy: 0.001 mA

34

Data unit: mA

--example:

Set the ACW current lower limit in STEP 1 to: 1mA

Setting command: FUNC: STEP: <num>: AC: LOWC 1

Query command: FUNC: STEP: <num>: AC: LOWC?

Return value: 1.000.

FUNC: SOURce: STEP: AC: TTIM

--format:

Set / Query ACW test time

Setting format: FUNC: STEP: <num>: AC: TTIM <time value>

FUNC: STEP: <num>: AC: TTIM? Query Format:

--Data <time value>:

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

--example:

Set the ACW test time in STEP 1 to: 1s

Setting command: FUNC: STEP: <num>: AC: TTIM 1

Query command: FUNC: STEP: <num>: AC: TTIM?

Return value: 1.

FUNC: SOURce: STEP: AC: RTIM

--format:

Set / Query ACW Rise Time

Setting format: FUNC: STEP: <num>: AC: RTIM <time value>

FUNC: STEP: <num>: AC: RTIM? Query Format:

--Data <time value>:

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

--example:

Set the rise time of ACW in STEP 1 to: 1s

Setting command: FUNC: STEP: <num>: AC: RTIM 1

Query command: FUNC: STEP: <num>: AC: RTIM?

Return value: 1.000.

FUNC: SOURce: STEP: AC: FTIM

Set / Query ACW Fall Time

--format:

Setting format: FUNC: STEP: <num>: AC: FTIM <time value>

Query format: FUNC: STEP: <num>: AC: FTIM?

--example:

Set the ACW fall time in STEP 1 to: 1s

Setting command: FUNC: STEP: <num>: AC: FTIM 1

Query command: FUNC: STEP: <num>: AC: FTIM?

Return value: 1.

FUNC: SOURce: STEP: AC: ARC

--format:

Set / Query ACW ARC Current Limit

Format:

FUNC: STEP: <num>: AC: ARC <current value>

FUNC: STEP: <num>: AC: ARC?

Query format:

--Data <current value>:

Data type: floating point

Data range: 0 ~ 20.0 mA (where 0 is OFF) RK9920

Data accuracy: 0.1 mA

Data unit: mA

--example:

Set the ARC current upper limit of ACW in STEP 1 to: 1mA

Setting command: FUNC: STEP: <num>: AC: ARC 1

Query command: FUNC: STEP: <num>: AC: ARC?

Return value: 1.000

FUNC: SOURce: STEP: AC: FREQ

--format:

Set / Query ACW test frequency

Setting format: FUNC: STEP: <num>: AC: FREQ <frequency>

Query format: FUNC: STEP: <num>: AC: FREQ?

--Data <frequency>:

Data type: Character

Data range: 50/60

Data accuracy:

Data unit: Hz

--example:

Set the ACW test frequency in STEP 1 to: 50Hz

Setting command: FUNC: STEP: <num>: AC: FREQ: 50

Query command: FUNC: STEP: <num>: AC: FREQ?

Return value: 50.

5.3.4 DC Setup Function Command Set

Note: The basic format refers to the AC Setup function command set.

FUNC: STEP: <num>: DC: VOLT

--Data <voltage value>:

Data type: floating point

Data range: 0.050 ~ 6.000

Data accuracy: 1

Set / Query the DCW voltage

Data unit: KV

36

FUNC: STEP: <num>: DC: UPPC

--Data <current value>:

Set / Query DCW upper limit current

Data type: floating point

Data range: 0.001 ~ 10.000 mA RK9920

RK9910 0.001 ~ 5.000 mA

Data accuracy: 0.001 mA

Data unit: mA

FUNC: STEP: <num>: DC: LOWC

--Data <current value>:

Data type: floating point

Set / Query DCW lower limit current

Data range: 0 ~ 10.0 mA (where 0 is OFF)

0 ~ 5.0 mA (where 0 is OFF)

Data accuracy: 0.001 mA

Data unit: mA

RK9920

RK9910

FUNC: STEP: <num>: DC: TTIM

--Data <time value>:

Set / Query DCW test time

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

FUNC: STEP: <num>: DC: RTIM, set / query DCW rise time

--Data <time value>

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

FUNC: STEP: <num>: DC: FTIM, set / query DCW fall time

--Data <time value>:

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

FUNC: STEP: <num>: DC: ARC, Set / Query DCW ARC current upper limit

--Data <current value>:

Data type: floating point

Data range: 0 ~ 20.0 mA (where 0 is OFF)

0 ~ 10.0 mA (where 0 is OFF)

Rk9920                       37

Rk9910

Data accuracy: 0.1 mA

Data unit: mA

FUNC: STEP: <num>: DC: RAMP

--format:

Set / Query DCW boost status

Setting format: FUNC: STEP: <num>: DC: RAMP: <ON / OFF> or <1/0>

FUNC: STEP: <num>: DC: RAMP :?

Query format:

--Data <voltage>:

Data type: Character

Data range: OFF (0), ON (1)

Data accuracy:

Data unit:

--example:

Set the RAMP status of the DCW in STEP 1 to: ON

Setting command: FUNC: STEP: <num>: DC: RAMP ON

Query command: FUNC: STEP: <num>: DC: RAMP?

Return value: ON.

5.3.5 IR Setup Function Command Set

Note: The basic format refers to the AC Setup function command set.

FUNC: STEP: <num>: IR: VOLT, set / query IR voltage

--Data <voltage value>:

Data type: floating point

data range:

0.050 ~ 1.000

Data accuracy: 1

Data unit: K V

FUNC: STEP: <num>: IR: UPPC, set / query the upper limit of IR resistance

--Data <resistance value>:

Data type: floating point

Data range: 0 ~ 1E4 (0 is OFF) MΩ

Data accuracy: 0.1 MΩ

Data unit: MΩ

FUNC: STEP: <num>: IR: LOWC Set / Query IR resistance lower limit

--Data <resistance value>:

Data type: floating point

Data range: 0.1 ~ 1E4 MΩ

Data accuracy: 0.1 MΩ

Data unit: MΩ

FUNC: STEP: <num>: IR: TTIM, set / query IR test time

--Data <time value>:

Data type: floating point                38

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

FUNC: STEP: <num>: IR: RTIM, set / query IR rise time

--Data <time value>:

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

FUNC: STEP: <num>: IR: FTIM, set / query IR fall time

--Data <time value>:

Data type: floating point

Data range: 0 ~ 999.9 (where 0 is OFF)

Data accuracy: 0.1

Data unit: s

FUNC: STEP: <num>: IR: RANG, set / query IR range

--Data <range value>:

Data type: integer

Data range: 0 ~ 5 (where 0 is AUTO, 1 is 10mA, 2 is 2mA, 3 is 200uA,

4 is 20uA, 5 is 2uA,)

5.4 SYSTEM subsystem command set

Command tree:

39

SYSTEM: DELAY, set / query startup delay.

--format:

Format: SYST: DELAy <time value>

Query format: SYST: DELAy?

--data:

Data type: float

Data range: 0.1 ~ 99.9

Data accuracy: 0.1

Data unit: s

--example:

Set DELAY to 1.0s

Setting command :: SYST: DELAY 1

--returned messages

Query command: SYST: DELAY ?, return value: START DELAY setting value, such as 1.000.

SYSTEM: STEP Set / Query the interval time of STEP.

--format:

Format: SYST: STEP <time value>

Query format: SYST: STEP?

--data:

Data type: float

Data range: 0.1 ~ 99.9

Data accuracy: 0.1

Data unit: s

--example:

Set STEPHOLD to 1.0s

Set command :: SYST: STEP 1

--returned messages

Query command: SYST: STEP ?, return value: STEPHOLD setting value, such as 1.000

SYSTem: GFI

Set / Query the status of GFI.

--format:

Setting format: SYST: GFI <ON / OFF> or <1/0>

Query format: SYST: GFI?

--Data <ON / OFF>:

Data type: Character

Data range: 0 (OFF), 1 (ON)

--example:

Set GR CONT to ON

Setting command: SYST: GFI ON or: SYST: GFI: 1

--returned messages

Query command: SYST: GFI ?, return value: status of GR CONT, such as 1.                        40

SYSTem: FAIL, set / query the status of AFTR FAIL.

--format:

Setting format: SYST: FAIL <0/1/2/3>

Query format: SYST: FAIL?

--Data <STOP / CONT / REST / NEXT>:

Data type: Character

Data range: 0 ~ 3

--example:

Set AFTR FAIL to STOP

Set command: SYST: FAIL 0

--returned messages

Query command: SYST: FAIL ?, return value: status of AFTR FAIL, such as 0.

SYSTem: LANGuage, set / query language status

--format:

Setting format: SYST: LANG <0/1>

Query format: SYST: GFI?

--Data <ON / OFF>:

Data type: Character

Data range: 0 (Chinese), 1 (English)

--example:

Set LANG to 0 (Chinese)

Set command: SYST: LANG 0

--returned messages

Query command: SYST: LANG ?, the return value: the status of LANG, such as 0.

SYSTem: PBEE / FBEE / KBEE settings / Query test pass / fail / buzzer status

--format:

Setting format: SYST: PBEE <ON / OFF> or <1/0>

41

Query format: SYST: PBEE?

--Data: <OFF / ON>

Data type: Character

Data range: 0 (OFF), 1 (ON)

--example:

Set BEEP to 1

Set command: SYST: BEEP 1

--returned messages

Query command: SYST: BEEP ?, Return value: Buzzer status, such as 1.

SYSTem: OFFSet Set / Query the status of the clear setting.

--format:

Setting format: SYST: OFFS <ON / OFF / GET> or <1/0 / GET>

Query Format: SYST: OFFS?

--Data <ON / OFF / GET>:

Data type: Character

Data range: 0 (OFF), 1 (ON), GET (get the clear value)

--example:

Set OFFS to ON

Setting command: SYST: OFFS ON or: SYST: OFFS 1

--returned messages

Query command: SYST: OFFS ?, return value: OFFS status, such as 1.

--Get the current value directly

Set command: SYST: OFFS GET

SYSTem: REset restores all settings to the default state.

--format:

Format: SYST: RES

42

5.5 MMEM subsystem command set

MMEM: SAVE saves the current file to a file number.

--format:

Format: MMEM: SAVE <file name>

--Data <file name>:

Data type: String

MMEM: LOAD Exports the file specified by the file number to the current.

--format:

Format: MMEM: LOAD <file name>

--Data <file name>:

Data type: String

5.6 FETCH Subsystem Command Set

FETCH is used to obtain the measurement results of the instrument.

--format:

Setting format: FETCh: AUTO <ON / OFF> or <1/0>

Query format: FETCh?

--Data <ON / OFF> or <1/0>

Data type: Character

Data range: 0 (OFF), 1 (ON)

--example:

Automatically set test data to ON

The command is: FETCh: AUTO ON or: FETCh: AUTO 1

--returned messages

Query command: FETCh ?, returns the current measurement result of the instrument.

Command syntax: FETCh?

After the instrument receives this command, the instrument will automatically send out the test results until the test is over.

Command syntax: FETCh?

After the instrument receives this command, the instrument will automatically send out the test results until the test is over.

Returned format:

Steps: Test items: test voltage (V), test current (mA), sorting results;

⑴ ⑵ ⑵ ⑶

⑴: Steps and test items, the separator between test items and data is (:).

⑵: The separator between test data is (,). The separator between different unit data is (;)

⑶: The separator between steps is (; + space). The data terminator defaults to (0x0A).

Note: 1. All data is in integer or floating point format, ASCII string.

2. By default, the data unit is the same as the FUNC setting instruction set.                              43

The test results are

STEP1: AC: 1000V, test current 1mA, result PASS.

STEP2: IR: 500V, test resistance 100M, result PASS.

Returned data format:

STEP1: AC: 1000, 1.000, PASS; (SPACE)

STEP2: IR: 500, 100.000, PASS; (0x0A)

5.7 Other control command set

* IDN query instrument model and version information

The query returns: <manufacturer>, <model>, <firmware> <NL ^ END>

Here: <manufacturer> gives the manufacturer name (ie REK)

<model> gives the model of the machine (eg RK9920 / 9910)

<firmware> gives the software version number (such as Version1.0.0)

For example: WrtCmd ("* IDN?");

| model | RK9920 series | RK9910 series |
| --- | --- | --- |
| basic skills | screen size | 5 inch TFT LCD screen |
|  | number key | Parameter setting digital input |
|  | Coding switch | Parameter selection and confirmation function |
|  | Up, down, left and right function keys | Parameter setting up and down selection function |
|  | LOCK keyboard lock function | Prevent accidental modification of test conditions or prohibit test conditions from being modified |
|  | Alarm function | Audible alarm |
|  | Communication Interface | RS232C、RS484 |
|  | USB | Copy, copy, store function |
|  | Control interface | HANDLER(PLC)、SINGAL |
|  | AC | voltage range | 0.050KV-5.000KV |
|  |  | Voltage waveform | Sine wave |

44

| Withstand voltage test | Distortion | ≤1% |
| --- | --- | --- |
|  |  | working frequency | 50Hz or 60Hz |
|  |  | Frequency accuracy | ±1.0% |
|  |  | Output Power | 100VA(5.000KV 20mA) | 50VA(5.000KV 10mA) |
|  |  | Voltage regulation | ± (1.0% + 50V) (rated power) |
|  | DC | voltage range | 0.050KV-6.000KV |
|  |  | Signal source frequency | 600Hz |
|  |  | Output Power | 50VA(5.000KV 10mA) | 25VA(5.000KV 5mA) |
|  |  | Voltage regulation | ± (1.0% + 100V) (rated power) |
|  | Voltage resolution | 1V |
|  | Voltage output accuracy | ± (2.0% setting + 5V) no load |
|  | Voltage test accuracy | ± (2.0% setting + 2V) |
|  | Voltage generation method | DDS signal source + AB power amplifier |
|  | Current test range | AC 0.001mA-20mA | AC 0.001mA-10mA |
|  |  | DC 0.1uA-10mA | DC 0.1uA-5mA |
|  | Current accuracy | ± (2.0% of reading + 5 words) | ± (2.0% of reading + 5 words) |
|  |  | ± (2.0% of reading + 5 words) | ± (2.0% of reading + 5 words) |
| Insulation resistance test | output voltage | 0.05KV-1.000KV |
|  | Voltage resolution | 1V |
|  | Voltage test accuracy | ± (2.0% of reading + 2V) |
|  | Output current | 10mA | 5mA |
|  | Output power | 10VA(1000V/10mA) | 5VA(1000V/5mA) |
|  | Ripple (1KV) | ≤3% (1KV no-load) | ≤3% (1KV no-load) |
|  | Resistance measurement range | 0.1MΩ-10GΩ | 0.2MΩ-10GΩ |

45

|  | Resistance display range (1000V) | 10mA(0.1MΩ-0.5MΩ) | 5mA(0.2MΩ-1MΩ) |
| --- | --- | --- | --- |
|  |  | 2 mA（0.5MΩ-5MΩ） | 1mA（1MΩ-10MΩ） |
|  |  | 200uA（5MΩ-50MΩ） | 100uA（10MΩ-100MΩ） |
|  |  | 20uA（50MΩ-500MΩ） | 10uA（100MΩ-1GΩ） |
|  |  | 2uA（500MΩ-10GΩ） | 1uA（1GΩ-10GΩ） |
|  | Resistance measurement accuracy | ≥500V | ≥500V |
|  |  | 1MΩ-1GΩ ± (5% of reading + 5 words) | 1MΩ-1GΩ ± (5% of reading + 5 words) |
|  |  | 1GΩ-10GΩ ± (10% of reading + 5 words) | 1GΩ-10GΩ ± (10% of reading + 5 words) |
|  |  | ＜500V | ＜500V |
|  |  | 0.1MΩ-1GΩ ± (10% of reading + 5 words) | 0.2MΩ-1GΩ ± (10% of reading + 5 words) |
|  |  | 1GΩ-10GΩ No precision required | 1GΩ-10GΩ No precision required |
|  | Current measurement accuracy | ± (1.5% of reading + 5 words) | ± (1.5% of reading + 5 words) |
| Arc Detect | Measuring range | AC 1mA-20mA |
|  |  | DC 1uA-20mA |
| Comparators |
| Judgment method | Window comparison |
|  | I down ON: When I down <IX <I, PASS; when IX ≤ I or IX ≥ I, FAIL (<I under condition I) |
|  | I down OFF: when IX <I, I pass, when IX≥I, FAIL, insulation resistance judgment method is the same as above. |
| Current cap setting I | AC | 0.001mA-20mA | 0.001mA-10mA |
|  | DC | 0.1uA-10mA | 0.1uA-10mA |
| Current upper limit setting I (LOWER OFF) | AC | 0.001mA-20mA | 0.001mA-10mA |
|  |  | 0.1uA-10mA | 0.1uA-10mA |
| Resistance upper limit setting | OFF-0.1MΩ-10GΩ | OFF-0.2MΩ-10GΩ |
| Resistance lower limit setting | 0.1MΩ-10GΩ | 0.2MΩ-10GΩ |

46

| Discrimination output | PASS / FAIL LCD and LED display separately, audible alarm |
| --- | --- |
| Parameter setting |
| Voltage rise time | 0.1s-999.9s |
| Voltage drop time | 0s-999.9s (only after withstand voltage PASS) |
| Voltage waiting time | 0.3s-999.9s (only DC withstand voltage, and meets the rise time + test time> waiting time) |
| Test time setting | 0.3s-999.9s (at TIMER ON time) |
| Time accuracy | ± (0.2% of setting value ± 0.1s) |

6.2 General technical indicators

| General technical indicators |
| --- |
| Working temperature and humidity | 0℃-40℃，≦75%RH |
| power supply | 100V-121V，198V-242V，47.5-63Hz |
| Power consumption | RK9920/A/B | ≦400VA | RK9910/A/B ≦300VA |
| Shape volume | 430mm×105mm×350mm |
| weight | 17KG | 15KG |

47

7, the instrument system upgrade steps:

1. Connect the instrument USB to the computer, the computer resource manager displays the U disk, copy the upgrade file to the root directory, and restart

Start the instrument to complete the upgrade. If there is an error during the upgrade, please contact us in time.

2. Press and hold STOP + STAR to restart the power supply, clear the data error caused by the version change, and restore the setting data to the default factory data.

Appendix: Parameter table

| Parameter / model | Program-controlled insulation withstand voltage tester | Programmable withstand voltage tester |
| --- | --- | --- |
|  | RK9920 | RK9910 | RK9920A | RK9910A | RK9920b | RK9910B |
| Withstand voltage test | Output voltage (KV) | AC:0.05-5.00 DC:0.05-6.00 | AC:0.05-5.00 DC:0.05-6.00 | AC:0.05-5.00 |
|  | Test accuracy | ± (2.0% setting + 2V) |
|  | Output accuracy | ± (2.0% setting + 5V) No load |
|  | Test current (MA) | AC:0.001mA-20mA DC:0.1uA-10mA | AC:0.001mA-10mA DC:0.1uA-5mA | AC:0.001mA-20mA DC:0.1uA-10mA | AC:0.001mA-10mA DC:0.1uA-5mA | AC:0.001mA-20mA | AC:0.001mA-10mA |
|  | Test accuracy | ± (2.0% of reading + 5 words) |
| Insulation test | The output voltage (KV) | DC:0.050-1.00 | / | / | / | / |
|  | Test accuracy | ±(2.0% setting + 2V) | / | / | / | / |
|  | Test resistance | 0.1MΩ-10GΩ | 0.2MΩ-10GΩ | / | / | / | / |
|  | Test accuracy | 1.0MΩ-10GΩ ±20% | 0.1MΩ-999MΩ ±10% | / | / | / | / |
|  | Discharge function | Automatic discharge after test | / | / | / | / |
| Arc detection |  |  |  |  |  |
| Measuring range | Corresponding current | 1mA-20mA |
| testing time | 0.1S-999.9S |
| Output frequency | 50Hz/60Hz |
| Input characteristics | 115V/230V±10% 50Hz/60Hz |
| Test alarm | Buzzer, LCD, FAIL indicator |
| Keyboard lock | Independent keyboard lock key |
| screen size | 5 inch TFT LCD |
| Communication Interface | HANDLER, RS232, RS485, USBDEV (computer interface), USBHOST (U disk interface) |
| Voltage rise time | 0.1S-999.9S |
| Test time setting (AC / DC) | 0.2S-999.9S |
| Voltage drop time | 0.1S-999.9S |
| waiting time | 0.2S-999.9S |
| Memory | 16M Flash, 50 test steps per file |
| size（W×H×D | 430mm×105mm×350mm |
| weight | 17kg | 15kg | 17kg | 15kg | 17kg | 15kg |
| Accessories | Test lead, ground lead, power lead |

48

8.RK99 series program-controlled medical pressure tester overview and parameters

This series of program-controlled withstand voltage tester is a high-performance safety tester designed with a high-speed MCU and large-scale digital circuits. The output voltage is increased and decreased. The frequency safety of the output voltage is controlled by the MCU, which can display the breakdown current and voltage value in real time, and has a software calibration function. It is equipped with a PLC interface, RS232C, RS485, USB device, USB Host interface, which can be easily combined with a computer or PLC to form a comprehensive test system. . It can quickly and accurately perform comprehensive safety measurement on household appliances, instruments, lighting appliances, electric heating appliances, computers, and information machines.

This instrument complies with the safety of IEC60335-1, GB4706.1, UL60335-1 household and similar electrical appliances Part 1: General requirements IEC60335-1, GB4706-1, UL60335-1 Information technology equipment is applicable to UL60065, GB8898, IEC60065 , Video and similar electronic equipment safety requirements IEC61010-1, GB4793.1 Safety requirements for electrical equipment for measurement, control and laboratory use Part 1: General requirements.

Application area

Components: diodes, triodes, high-voltage silicon stacks, various electronic transformers, connectors, high-voltage electrical appliances

Household appliances: TV, refrigerator, air conditioner, washing machine, dehumidifier, electric blanket, charger, etc.

Insulation materials: heat shrinkable sleeves, capacitor films, high voltage sleeves, insulating paper, insulating shoes, insulating rubber gloves, PCB circuit boards, etc.

Instrumentation: oscilloscope, signal generator, DC power supply, switching power supply and other complete machines

Lighting appliances: ballasts, street lights, stage lights, portable lights, etc.

Electric heating appliances: electric drill, pistol drill, gas cutter, grinder, grinder, electric welding machine, etc.

Wires and cables: high-voltage lines, closed cables, cables, silicone rubber cables, etc.

Performance characteristics

1. AC / DC withstand voltage function adopts DDS digital signal synthesis technology to produce accurate, stable, pure and low distortion waveform

2, adjustable high-voltage rise and fall time, to meet the requirements of different test objects with arc detection function test results can be saved simultaneously

3, with dual frequency comprehensive test, frequency range 50Hz, 60Hz user-friendly operation interface, support for digital key input, dial input, simpler operation

4. Complete operation help tips, which can effectively improve user efficiency, support character file name input, the maximum length of the file name is 12 characters

5. The test steps and system status information are displayed synchronously, which makes it easy to understand the details of the test steps and the system status during the test.                               49

6, Chinese and English bilingual operation interface, adapt to the needs of different users, support large-capacity storage. Adapt to different test application requirements

Normal pass     Slight arc    Severe arc

Appendix: Parameter table

| Parameter / model | Program-controlled insulation withstand voltage tester | Programmable withstand voltage tester |
| --- | --- | --- |
|  | RK9920 | RK9910 | RK9920A | RK9910A | RK9920b | RK9910B |
| Withstand voltage test | Output voltage (KV) | AC:0.05-5.00 DC:0.05-6.00 | AC:0.05-5.00 DC:0.05-6.00 | AC:0.05-5.00 |
|  | Test accuracy | ± (2.0% setting + 2V) |
|  | Output accuracy | ± (2.0% setting + 5V) No load |
|  | Test current (MA) | AC:0.001mA-20mA DC:0.1uA-10mA | AC:0.001mA-10mA DC:0.1uA-5mA | AC:0.001mA-20mA DC:0.1uA-10mA | AC:0.001mA-10mA DC:0.1uA-5mA | AC:0.001mA-20mA | AC:0.001mA-10mA |
|  | Test accuracy | ± (2.0% of reading + 5 words) |
| Insulation test | The output voltage (KV) | DC:0.050-1.00 | / | / | / | / |
|  | Test accuracy | ±(2.0% setting + 2V) | / | / | / | / |
|  | Test resistance | 0.1MΩ-10GΩ | 0.2MΩ-10GΩ | / | / | / | / |
|  | Test accuracy | 1.0MΩ-10GΩ ±20% | 0.1MΩ-999MΩ ±10% | / | / | / | / |
|  | Discharge function | Automatic discharge after test | / | / | / | / |
| Arc detection |  |  |  |  |  |
| Measuring range | Corresponding current | 1mA-20mA |
| testing time | 0.1S-999.9S |
| Output frequency | 50Hz/60Hz |
| Input characteristics | 115V/230V±10% 50Hz/60Hz |
| Test alarm | Buzzer, LCD, FAIL indicator |
| Keyboard lock | Independent keyboard lock key |
| screen size | 5 inch TFT LCD |
| Communication Interface | HANDLER, RS232, RS485, USBDEV (computer interface), USBHOST (U disk interface) |
| Voltage rise time | 0.1S-999.9S |
| Test time setting (AC / DC) | 0.2S-999.9S |
| Voltage drop time | 0.1S-999.9S |
| waiting time | 0.2S-999.9S |
| Memory | 16M Flash, 50 test steps per file |
| size（W×H×D | 430mm×105mm×350mm |
| weight | 17kg | 15kg | 17kg | 15kg | 17kg | 15kg |
| Accessories | Test lead, ground lead, power lead |

10. Random attachments

RK26003A withstand voltage test clip * 1

RK26003B withstand voltage grounding clamp * 1

RK-8G-1 Cross uncontrolled high voltage probe * 1

Power cord 1.8m * 1

BNC line double-headed line * 2 (RK99 program-controlled medical pressure withstand instrument series)

Manual * 1

# Text boxes

4.4.1 Start test

After the instrument is in the test mode, check the test conditions and the DUT is connected correctly, and then press the START key to start the test.

4.4.2 Delay Test Delay

After starting the test, the delay before the first step is delayed according to the set delay time in the system interface. The delay between multiple steps is maintained according to the items in the system interface.

A

100

10

C

B

STEP

SYSTEM

DELAY

PBEE

SHORT

OFFSET

SAFE

GFI

FALL

FBEE

LANGUAGE

TIME

DATE

RES