SHENZHEN MEIRUIKE ELECTRONIC TECHNOLOGY CO.,LTD

# Contents

Chapter 1 Overview	1.2

Chapter 2 Specifications	2.3.4

Chapter 3 Function Description	5

Chapter 4 Operation Manual	6.7.8.9

Chapter 5 Calibration Inspection and maintenance9.10.11

## Chapter 1 Overview

  - Product Summary

RK2681N/2681AN type Insulation Resistance Tester is a measurement which can measure the insulation of the electronic parts, dielectric material, and machines. It has the function of test quickly, good stability, easy operation and determine the eligibility of the product.

The instrument is correspond with the requirements and conditions of Group Ⅱ of Ministry of Electronics Industry Standards GB6587.1 «environmental testing of electronic measuring instruments master>>：

A：Temperature：0~40℃

B：Relative Humidity：<70% C：Atmospheric pressure：86~106kPa

  - Confirmed Accessories	before using

After receipt of this tester, please confirm the following items as shown in order to safeguard your interests:

A：Whether the	appearance of the product is good.

B：Standard accessories in the table below, make sure about them.

If there have some damage or lack of attachments， Please contact with us or the agents where you buy them so that we can give you the best service.

| Name | Quality | Notes |
| --- | --- | --- |
| Power Cable | 1 |  |
| Testing wire | 1 | Generally testing wire |
| Manual | 1 |  |
| Testing Transposon |  | OPTIONAL ACCESSORIES |

  - Environmental requirements

    - Operating temperature 0 ℃ ~ 40 ℃, relative humidity:≤ 70%. Do not let the tester place at dust, multi-vibration, as well as direct sunlight or corrosive gases environment.

    - In order to prevent this instrument internal temperature rises too much, please take the  instrument and the surrounding objects in more than 10 cm distance, and do not block the  right ventilation holes to protect the good of the instrument accuracy.

    - Although this instrument has been design for the interference with AC power supply.. But also please use it in the environment where the interference is little if possible.   Please install some power supply filtering device if the interference can’t be avoided.

    - Storage Temperature of the instrument is -25℃~+70℃.If the instrument will not be used for a long time. Please send it into the box and place it where no direct sunlight and dry place, in order to ensure accuracy.

  - Power input

Please make sure the power switch is in OFF before contact with the AC power, and make sure the voltage input is AC 220V±10%,50Hz or 60Hz.

  - Fuse

There is a fuse is at the back panel,please note when change it:

    - Please take the switch of the power in OFF and then cut off the power. 1．5．2．The Resistance of the fuse is 15Ω,if not, it is broken. Please don’t check it by your

eyes.

    - Please use a little Screwdriver to take the broken fuse out.

    - Replace the new fuse (as same as the broken one) and put it into the AC socket.

  - Warm-up Time

You can use all the functions of this tester after power is turned on. However, to  achieve the target within the accuracy, please warm more than 30 minutes.

# Chapter 2 Specifications

#### Test voltage：

RK2681N:

10VDC	Voltage magnification×0.1

25 VDC		Voltage magnification×0.25 50 VDC	Voltage magnification×0.5 100 VDC	Voltage magnification×1

250 VDC	Voltage magnification×2.5 500 VDC	Voltage magnification×5

RK2681AN:

10VDC			Voltage magnification×0.1 50VDC	Voltage magnification×0.5 100VDC		Voltage magnification×1 250VDC		Voltage magnification×2.5 500VDC		Voltage magnification×5 1000VDC			Voltage magnification×10

Accuracy:±2%(when the test port is turn off);the measured voltage of the load is over 95%

  - Test range

#### RK2681N：		100KΩ		~		5TΩ（105~5×1012Ω） 10V：	100 KΩ	~		100GΩ

25V：	250 KΩ	~	250GΩ

50V：	500 KΩ	~	500GΩ

100V：	1MΩ	~	1TΩ

250V：	2.5MΩ	~	2.5TΩ

500V：	5MΩ	~	5TΩ

RK 2681AN：		100KΩ	~		10TΩ（105~1013Ω） 10V：	100 KΩ	~		100GΩ

50V：	500 KΩ	~	500GΩ

100V：	1MΩ	~	1TΩ

250V：	2.5MΩ	~	2.5TΩ

500V：	5MΩ	~	5TΩ

1000V：	10 MΩ	~	10TΩ

  - Accuracy：

Impedance	<1GΩ	±3%±0.5 digits.

Impedance	>=1GΩ	±5%±0.5 digits.

Impedance >100GΩ	±10%±0.5 digits.

  - Charge Time

Electronic machinery, media materials, resistive elements ： Charge Time

<0.1seconds.

Capacitor：Charge Time 0.5～10 seconds.

  - Preset range

（1.0～9.9）×Voltage magnification×Rate Switch (MΩ),Accuracy：±0.3%±0.5 digits.

  - Judge and select

This instrument can judge and select，you can set a value with the preset switch first. If the testing value is less than it, the light will bright, and the buzzer will ring. It means that the parts be measured is not good. And if the testing value is over than the preset value, it means it is good.

  - Panel instructions

When	button has not been press down ，the display of the panel is Preset values.

When	button has been press down，the display of the panel is the measured values.

  - Discharge

When the witch is turn to 放 电 ，it will discharge to the measured parts， Resistance of discharge is depended on the testing voltage. There is no output voltage at the test port。

  - User guide：

    - Temperature 0～40℃，Relative Humidity：≤70% 。2．9．2．Storage Environment：-25℃～70℃。2．9．3．Power input：220V±5%。

Frequency：50Hz/60Hz。Power：≤30V

  - Volume and weight 2．10．1．Volume：360mm（L）×360mm（W）×135mm（H）

2．10．2. Weight： 7Kg

Chapter 3 Function Description

3．1．Panel：RK2681N/2681ANtypeInsulation Resistance meter（100μA panel）。

3．2．  switch：there are six times stalls:×100，×101，×102，× 103，×104，×105。

3．3．   switch：when you put the switch in，the instrument discharge the tested parts.

3．4．  switch：there are six voltage stalls：

RK2681N：10V/25V/50V/100V/250V/500V RK2681AN：10V/50V/100V/250V/500V/1000V

3．5．  switch：The switch of the panel state；

When turn it to setting，the display is Setting value；     When turn it to testing,，the display is measured value.

  - Dial setting

Composed by two 8421 to sort the resistance.

  - Voltage output port:

Negative high voltage output，which connect with one port of the testing wire. Please be careful with it because there have high voltage.

  - Testing port：the port with connect with another port of the testing wire. 3．9．“∞”Adjustment potentiometer

“∞”Adjust the potentiometer and the magnification ranges to make the pointer of the meter to “∞”.

  - Ground port.

  - When the light is well bright, it means the resistance of the measured element is less than the preset. And the measured element is not good.

When the light is not bright, it means the resistance of the measured element is over the preset. And the measured element is good.

  - Power switch. 3．13．Model, Name, Brand

Model：RK2681N or RK2681AN Name：Insulation Resistance Tester Brand：Rek

# Chapter 4 Operation Manual

Please read this manual carefully first to know the test method and the basic principles.

  - Test Method:

  - You must to know before Setting the Test parameters. 4．2．1

### Equivalent circuit

| Magnification | Impedance Input |
| --- | --- |
| 100 | 1KΩ |
| 101 | 10 KΩ |
| 102 | 100 KΩ |
| 103 | 1 MΩ |
| 104 | 10 MΩ |
| 105 | 100 MΩ |

    - The figure shows the relationship between the input impedance and the magnification of the equivalent circuit. The test ranges of the measured part or machine and the relationship of the voltage and the magnification the following table:

    - Adjust every potentiometers of the instrument，the pointer should pointer to the“∞”always.

When the range is over ×105：

      - When there is no testing fixer or the testing cable，the pointer of the panel should point to “∞”；

b．When the test port has been connected the testing fixture or the testing cable, and have not been connected with the measured part Rx, the pointer of the panel is allow to have a drift. Because the resistance at this time is about 100MΩ.It said that this testing cable just like a "antenna" which can receive the signal outside.

When contact with Rx，there will be no interference signal disappears. And it will not affect the accuracy.

### Rk 2681N:	UnitsΩ

| 倍率范 围 电 压 |  | 0 |  | 1 |  | 2 |  | 3 |  | 4 |  | 5 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 10 |  | 10 |  | 10 |  | 10 |  | 10 |  | 10 |  |
| 10V | 100K～1M | 1M～10M | 10M～100M | 100M～1G | 1G～10G | 10G～100G |
| 25V | 250K～2.5M | 2.5M～25M | 25M～250M | 250M～2.5G | 2.5G～25G | 25G～250G |
| 50V | 500K～5M | 5M～50M | 50M～500M | 500M～5G | 5G～50G | 50G～500G |
| 100V | 1M～10M | 10M～100M | 100M～1G | 1G～10G | 10G～100G | 100G～1T |
| 250V | 2.5M～25M | 25M～250M | 250M～2.5G | 2.5G～25G | 25G～250G | 250G～2.5T |
| 500V | 5M～50M | 50M～500M | 500M～5G | 5G～50G | 50G～500G | 500G～5T |

Rk 2681AN:	Units Ω

| 倍率范 围 电 压 |  | 0 |  | 1 |  | 2 |  | 3 |  | 4 |  | 5 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 10 |  | 10 |  | 10 |  | 10 |  | 10 |  | 10 |  |
| 10V | 100K～1M | 1M～10M | 10M～100M | 100M～1G | 1G～10G | 10G～100G |
| 50V | 500K～5M | 5M～50M | 50M～500M | 500M～5G | 5G～50G | 50G～500G |
| 100V | 1M～10M | 10M～100M | 100M～1G | 1G～10G | 10G～100G | 100G～1T |
| 250V | 2.5M～25M | 25M～250M | 250M～2.5G | 2.5G～25G | 25G～250G | 250G～2.5T |
| 500V | 5M～50M | 50M～500M | 500M～5G | 5G～50G | 50G～500G | 500G～5T |
| 1000V | 10M～100M | 100M～1G | 1G～10G | 10G～100G | 100G～1T | 1T～10T |

M:106	G:109	T:1012

### The parameter about Voltage, magnification

8

  - Test parameter setting 4．3．1．Voltage Setting

Place the switch of the voltage band on the corresponding voltage profile, accordance with the relevant technical requirements. Please note at that the testing fixture has been put on electricity. Don't touch it to avoid an electric shock.

4．3．2．Magnification Setting

Please set the magnification according to the 4．2．2 4．3．3．Maximum setting

According to different technical requirement of the products set different values with the dial switch, the point display on the panel must meet the value on the dial switch. If the value be set is 2．0,the point display on the panel should be 2．0,but the lower limited insulation resistance value should not be 2．0.For example, the voltage you set is 250V，the magnification is ×102，the value on the dial switch is 2.0,then the lower limit set value is (2.0×102×2.5)MΩ=500 MΩ

    - Charge time setting

      - When the part which be tested is a pure resistance, it does not need to be charge. You can read the insulation resistance directly.

      - When the part which be tested is a pure capacitance, the charge time is not only depended on the capacitance and its quality, but also , but also be affected by the speed of testing.

c．The discharge	The current be discharged is constant current ，the speed of discharge is depend on the voltage of the testing port and the capacity of itself， Jcharge=KCU。

  - Examples

If there have a manufacturer need to sort a group of metalized film capacitors, the model is CBB-250V-2.2μF, insulation resistance Rx> 100G.

    - Turn the Power switch at ON, connected to power, the power indicator light, you can test after preheating for 15 minutes.

    - Adjust the voltage range to the 250V,then the voltage output is 250V, Take the	magnification switch to×104,According to 4.2.2 the table.

If the lower limited value at 100GΩ,the data you must set is : 100×109/(2.5×104×106)=4.0

    - Set the appropriate charging time: T=KCU If K=5000,C=2.2μF,U=250V

Then T=5000×2.2×106=2.75 seconds, the charge time is about 3 seconds.

    - You can test with the testing fixer or the testing cable. 4．4．5．Discharge to the measured capacitance, Please do that out of the machine.

Because when the testing voltage is 250V, the resistor of the discharge in the machine is about 2.5KΩ, it needs some time to discharge.

Chapter 5 Calibration Inspection and maintenance

  - Generation and been testing of DC Power supply 5．1．1．Generation of the 10 V DC Power

It generate by resistance R9 , R18 and PotentiometerW1，adjust Potentiometer W1 you can adjust the accuracy.

    - Generation of the other DC Voltage（RK2681AN,for example）

The voltage generated by V20 regulator tube is 5V.And the voltage generated by N1 is 0V.The current generated by R17 and W2 is constant current. Adjust W2，let it generate1mA constant current. The voltage generated by R51、R52、R53、R54、R55、R56 、 R57 、 R58 ， are different voltages. For example,R51generate50Vvoltage ， R52generate100Vvoltage	，	R53	and	R54generate250Vvoltage	，	R55	and R56generate500Vvoltage，R57 and R58 generate 1000V voltage.

    - If there is too much error of the voltage or there is no voltage, please check each circuit elements of the 5.1.2, and check is there damage about the circuit elements N1, V27, V28, V29 and peripheral resistance of diode, check and replacement. 5．1．4．Please check the connection with the following figure. Its test should meet to 2.1, if not satisfied, check the components and accuracy of the components

with5.1.1-5.1.3.

  - The testing of the Max charge current

    - The charge current is constant current when the tester discharge to the parts, the current I=20mA.It depend on the circuit N2，R44，R25，R2 所，I=R2*Vcc/

（R12*R44）.

    - The testing of the Max discharge following figure，put the voltage switch at 100V，put on the switch K less than 10 seconds ,turn off K after testing. If the charge current is error too much, please check it according to 5.2.1.You can also amend the charge current according to 5.2.1, you just need to amend the ratio of the R2 and R25 resistor.

5．3．1．Connect the standard resistor at the testing port and test it according to the following table. The error should meet to the 2.3 requirements.

| Test voltage | Standard Resistance | Measured resistance |
| --- | --- | --- |
| 100V | 1M |  |
|  | 10M |  |
|  | 100M |  |
|  | 1G |  |
|  | 10G |  |
|  | 100G |  |

| Standard Resistance | Test voltage | Measured resistance |
| --- | --- | --- |
| 100M | 10V |  |
|  | 25V |  |
|  | 50V |  |
|  | 100V |  |
|  | 250V |  |
|  | 500V |  |
|  | 1000V |  |

# Text boxes

Set/Measure

Set/Measure

!!!

—— Note! ——

- Please be careful while testing, there have high voltage  on the tested thing.

- If the component has polarity, please note the correct polarity connection.

- Isolated test environment and control the temperature and humidity to ensure the accuracy and stability.

- Don’t let the Shield of Test port （+）touch	to the shell of the instrument.

Rx

(+)

(-)

测试电压

Rx

外壳隔离

测试电压

Rk2681N/AN

6 1/2 Multimeter

Rx

COM

V

Rk2681N/AN

6 1/2 Multimeter

Rx

COM

I

K